## Cal
Calibration Sheet for SEAGLIDER |  | Serial No. | 139
Date: | 2009-07-01 18:19 |  | (master to all sheets)
Time: | 39995.7634
Software Revision | Rev 66.04/1579? - Nov 2008
Pitch Mass |  | Last updated | 2007-10-09
Comments
| Headroom | 25 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | cm travel*
Minimum (full forward) | 410 | 435 | -11.3553
Maximum (full aft) | 3749 | 3724 | 3.7741
Stroke Length |  |  | 15.1294
Pitch Center | (as defined by $C_PITCH) | 2903.5355 | 0
(15.32 cm full stroke)
Conversion constant | 0.0046 | cm per AD | ($PITCH_CNV)
(conversion inverse) | 217.39 | AD per cm
*equation | cm = (AD –center)* conversion factor (cm/AD)
(-) cm is fwd of center, (+) cm is aft of center
Roll Mass |  | Last updated | 2007-10-11
Comments
| Headroom | 150 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | degrees roll*
Full roll to port | 77 | 227 | -55.7767
Full roll to starboard | 4081 | 3931 | 48.9354
Dive Roll Center |  | 2200 | ($C_ROLL_DIVE)
Climb Roll Center |  | 2200 | ($C_ROLL_CLIMB)
Conversion constant | 0.02827 | degree per AD | ($ROLL_CONV)
(conversion inverse) | 35.37 | AD per degree
*Equation: | deg=(AD-center) * conversion factor
Variable Buoyancy Drive (VBD) |  | Last updated | 2007-06-28
Comments
LeDuc Pump SN | 032
Boost Pump SN | 1327473
| Headroom | 100 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | cm3 (cc) oil*
$VBD_Min (Ext Bladder full) | 350 | 450 | 514.0269
$VBD_Max (Ext Bladder empty) | 3960 | 3860 | -322.4325
Total movable oil volume |  |  | 836.4594
Neutral Trim (from Ballast sheet) | ($C_VBD for neutral) | 2545.537 | (nomially 0 at spec rho)
Conversion constant | -0.2453 | cc per AD | $VBD_CNV
(conversion inverse) | -4.0767 | AD per cc
*equation: | cc oil=(AD@neutral trim)* conv. factor (cc/counts)
(1 cc oil = 1.025 g buoyancy)
Pressure Sensor |  | Last updated | 2007-08-06
(all values fed from press-cal worksheet including date) |  | Comments
Values from Pres Cal sht
Manufacturer | Paine
Model # | 211-75-710-05
Serial # | 214721
Calibration date | 38423
A/D Gain | 128 |  | PSI = Slope *AD_Count + Y_Intercept
Atmos @ cal. | 14.7
Temp @ cal. | 0 | Psig/AD count
Calibrated Slope | 0.00011633 | (either PresCal-B49 or PreCal-I56)
Cal. Y-Intercept | -10.8728 | (Changed each time sensor set for “sealevel”)
Depth Offset | -7.4479 | Meters
| Conversion Factor | 0.685 psig/meter
Internal Pressure |  | Last updated | 2008-12-04
start | 14.70, RH 30.28 | end | 9.68, RH 29.0
Compass |  | Last updated | 2008-12-08
Manufactuer | Sparton
Model number | SP3003D
Serial Number | C143
Compass SW version # | 7.4
Compass carier PCB | PN 52281 rev 1.2
Sand Point Calibration Date | 2008-12-04
TCM2MAT file creation date | 2008-12-07
Main Board |  | Last updated | 2008-12-05
Comments
Model # | 52290 | (52290=normal glider)
Mainboard Revision | B.4
Mainboard Serial Number | 52290-032
Computer Type | TT8 | Onset Computers
Computer Serial Number | 1237116 | replaced 12/05/08
Disk system | CF8V2 | Persistor Corp.
Serial Number | 1989
Compact Flash Card Manu. | Silicon Systems
CF card size | 256 MB
CF card serial/ID # | SSD-C25MI-3005
20784
Watchdog (PIC) setting (min) | 10 min | (10min = dips 2&4 on)
OEM Board & Sub-systems |  | Last updated | 2007-08-06
Comments
OEM Board Serial Number | SFC-OEM-040
GPS Manufacturer | Garmin
GPS Model Number | 15H-W
GPS Serial Number | 69784881
GPSI Antenna SN | 059 | (fab by Ocean Eng Services)
DC-DC converter SN | 15749A621
Transducer Manu. | ITC
Transducer Model | 3013
Transducer SN | 4009
Acoustic Transponder |  | Last updated | 2008-12-05
Comments
Manufacturer | AAE
Model | 955
Main-board SN | 041
Sub-board SN | 291
AAE software ver # | ver 3.061
Interrogate Frequency | 12.5 | KHz
Reply Frequency | 13.5 | KHz
DIP switch settings | 9/D | (sw-1/sw-2)
Iridium Modem |  | Last updated | 2007-08-06
Comments
Model | 9522A
IMEI Number | 300224010901010
SIM Card Owner | Univ of Hawaii
SIM Card SN | 8988169414000843571
SBE C and T Sensors |  | Last updated | 2008-12-04
Comments
Calibration date | 2008-11-16
Temp PCB SN | 66
Cond PCB SN | 66
Conductivity Calibration Coefficients |  | Temperature Calibration Coefficients
g = | -9.8723 | g = | 0.0043402
h = | 1.0933 | h = | 0.00063043
i = | -0.001236 | i = | 2.4021e-05
j = | 0.00017717 | j = | 2.5358e-06
Cpcor (nom) = | -9.57e-08 | f0 = | 1000
Ctcor (nom) = | 3.25e-06
Optical Sensor(s) |  | Last updated | 2007-08-05
Manufacturer | WET Labs | Calibration date | 2008-06-01
Model | BB2F-VMG
Serial Number | 380
Dissolved Oxygen Sensor |  | Last updated | 2009-07-01
|  | Calibration date | 2009-01-06
Manufacturer | SeaBird | Sea-Bird 43f coefficients
Model | 43f | Soc = | 0.00023
Serial Number | 43f-0092 | Foffset= | -836
|  | Boc= | NA
|  | Tcor= | NA
|  | Pcor= | NA
|  | A= | -0.00025244
|  | B= | 3.5548e-05
|  | C= | -8.9129e-07
|  | E= | 0.036
|  | Tau20 | 3.61

## Trim
| SEAGLIDER - Trim and Balance |  | Date: | 2009-07-01 18:19
| SG | 139 | Time: | 39995.7634 | [7 empty] | Spreadsheet Usage Notes from Russ Light
Comments:
[13 empty] | Fwd Fairing assy scale wt: | 11026
[13 empty] | Aft fairing assy scale wt: | 4644.4
|  | Enter data only where font is RED | [10 empty] | Fairing screws scale wt: | 39
| Density Trim | [11 empty] | Pupa assy scale wt: | 36195
| Total Weight (in air), summed items | 52315.8 | gm | [9 empty] | Total glider wt by assy scale wts: | 51904.4
| Total Scale Weight | 51904 | gm | post PS by sum of assys on kirks scale 12/16
| Weight difference/"error" | -411.8 | gm
| Variable wt, including nose-wt-plates | 3475.2 | gm
| Displaced Volume, measured in tank | 50276.9526 | cc | diff | see Ballast worksheet for more volume details
| Displaced Volume, summed items | 50915.6204 | cc | 638.6678
| Water Density | 1.0275 | gm/cc
| Net Buoyancy (relative to neutral, positive floats, no thrust) | 0 | cc |  | predicted using this spreadsheet, actual may vary!
| Internal Oil Stroke (100% is all inside, ext. bladder empty) | 61.45% |  | 2545.537 | A/D counts | (input %, not counts)
[6 empty] | ($PITCH_VBD_SHIFT) | 0.00167 | cm per cc
Pitch & Roll Trim
| Longitudinal CG | 89.5445 | cm | 35.2538 | inches
| Longitudinal CB | 89.5367 | cm | 35.2507 | inches
| 4 | 0.0078629 | cm
| Pitch-mass Stroke (100% is fully aft, nominal is 70%) | 75.05% |  | 2903.5355 | A/D counts | (input %, not counts)
| Pitch-mass (batt) Location | 11.4983 | cm
| Vehicle Pitch Angle (positive is nose up) | 0.7202 | deg | (goal is change of 15-20 deg per cm travel, use goal-seek on Location by +/-1cm to test)
| Vertical CG (VCG) | -0.57384 | cm | (goal is ~ -0.56 from prior builds)
| Computed VCB | 0.05167 | cm | (goal is ~0.05 from prior builds)
| VCB-VCG Separation | 0.62551 | cm | (goal is ~0.61 from prior builds)
| Roll-mass (batt) Angle | 0 | deg
| Vehicle Roll Angle | #N/A | deg | (goal is 20 deg Vehicle roll per 40 deg Battery Roll, but eqn had circular reference, now repl w/"20")
[5 empty] | Total |  |  | Moment |  |  |  |  | Moment |  |  | Moment | Moment
Assy. |  | Part # | Qty | Weight | Weight | LCG | LCG | LCG | Volume | Total Vol | LCB | LCB | LCB | VCG&B | VCG&B | VCG | VCB
Assy. Level | Description | (or Assy Wt) |  | grams | grams | inch | cm | gm-cm | cc | cc | inch | cm | cc-cm | inch | cm | gm-cm | cc-cm
Complete fairing & antenna
| Subtotal (summed) Weight | 16086.5 | gm | grams: diff
| Scale Weight | 15670.4 | gm | grams: -416.1
| Summed volume | 8110.6916 | cc
|  |  |  | sum-diff's | -464
Complete pupa
| Subtotal (summed) Weight | 36242.9 | gm | grams: diff
| Scale Weight | 36195 | gm | grams: -47.9
| Summed volume | 42804.9289 | cc
|  | Total weights and volume - Flying SG - without adjustment |  | grams: 52368.4; gm-cm: 4675662.0757; cc: 50915.6204; cc-cm: 4558815.1658; gm-cm: -30020.9428; cc-cm: 2630.789
|  | Adjustment to reflect trims found in tank & Port Susan |  |  | -52.6 | -66.8533 | -169.8073 | 8931.8634 | 0 | 0 | 29.8601 | 75.8447 | 0 | 0 | 0 | 0 | 0
|  | Total weights and volume - Flying SG - adjusted, reported above |  | grams: 52315.8; gm-cm: 4684593.9391; cc: 50915.6204; cc-cm: 4558815.1658; gm-cm: -30020.9428; cc-cm: 2630.789
0 | Flying Seaglider | 49801
1 | (8x) 10-32 x 3/8 FHMS, joint ring to aft fairing | 55240 | 1 | 11.7 | 11.7 | 46 | 116.84 | 1367.028 | 1.4576 | 1.4576 | 46 | 116.84 | 170.3037 | 0 | 0 | 0 | 0
1 | (8x) 1/4-28 x 1/2 FHMS, Aft fairing to endcap | 55294 | 1 | 27.3 | 27.3 | 49.5 | 125.73 | 3432.429 | 3.401 | 3.401 | 49.5 | 125.73 | 427.6104 | 0 | 0 | 0 | 0
1 | Assy, Forward Fairing | 55494
| Subtotal Weight | 11015.2 | gm | grams: diff
| Scale Weight | 11026 | gm | grams: 10.8
2 | Forward Fairing | 49897 | 1 | 5600.1 | 5600.1 | 28.22 | 71.6788 | 401408.4479 | 3394.4305 | 3394.4305 | 28.22 | 71.6788 | 243308.7048 | 0 | 0 | 0 | 0
2 | Joint Ring, Fwd to Aft Fairing (bonded to fwd fairing before paint; Long ring = 875g, short ring 688g ) | 49834 | 1 | 694.7 | 694.7 | 45.85 | 116.459 | 80904.0673 | 256.0634 | 256.0634 | 45.85 | 116.459 | 29820.8873 | 0 | 0 | 0 | 0
2 | Nose weight (bonded to fwd fairing after paint) | 49844 | 1 | 1245.2 | 1245.2 | 3.715 | 9.4361 | 11749.8317 | 147.9739 | 147.9739 | 3.715 | 9.4361 | 1396.2961 | 0 | 0 | 0 | 0
2 | Nose weight plates | 55448 | 1 | 741.4 | 741.4 | 3.715 | 9.4361 | 6995.9245 | 88.1046 | 88.1046 | 3.715 | 9.4361 | 831.3636 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 2" SHCS + 1/4 LW brass, nose plate to nose base (4 - 5 plates) | 55186 / 55253 | 0 | 28.5 | 0 | 3.715 | 9.4361 | 0 | 3.5505 | 0 | 3.715 | 9.4361 | 0 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 1.5" SHCS + 1/4 LW brass, nose plate to nose base (2 -3 plates) | 55185 / 55253 | 1 | 23.5 | 23.5 | 3.715 | 9.4361 | 221.7483 | 2.9276 | 2.9276 | 3.715 | 9.4361 | 27.6253 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x .75" SHCS + 1/4 LW brass, nose plate to nose base (1 plate) | 55473 / 55253 | 0 | 15.5 | 0 | 3.715 | 9.4361 | 0 | 1.931 | 0 | 3.715 | 9.4361 | 0 | 0 | 0 | 0 | 0
2 | Trim Lead at ring - bottom (5" strip) | 52442 / 55487 | 1 | 1217.5 | 1217.5 | 41.7 | 105.918 | 128955.165 | 107.7815 | 107.7815 | 41.7 | 105.918 | 11416.0026 | -5.62 | -14.2748 | -17379.569 | -1538.5596
2 | Trim Lead at ring - port side (5" strip) | 52490 / 55488 | 1 | 233.4 | 233.4 | 41.7 | 105.918 | 24721.2612 | 20.6622 | 20.6622 | 41.7 | 105.918 | 2188.4969 | 0 | 0 | 0 | 0
2 | Trim Lead at ring - starboard side (5" strip) | 52490 / 55488 | 1 | 381.1 | 381.1 | 41.7 | 105.918 | 40365.3498 | 33.7376 | 33.7376 | 41.7 | 105.918 | 3573.4198 | 0 | 0 | 0 | 0
2 | Trim Lead at ring - top (5" strip) | 52490 / 55488 | 1 | 476.1 | 476.1 | 41.7 | 105.918 | 50427.5598 | 42.1477 | 42.1477 | 41.7 | 105.918 | 4464.1962 | 5.62 | 14.2748 | 6796.2323 | 601.6495
2 | Trim Lead tape at ring (2.3g per 5.75"pc dbl-stick tape, 2.6g per 6"pc duct tape) | 55298 / 55406 | 1 | 18.5 | 18.5 | 41.7 | 105.918 | 1959.483 | 20.5556 | 20.5556 | 41.7 | 105.918 | 2177.2033 | 0 | 0 | 0 | 0
2 | Trim Lead in middle - bottom (5" strip) | 52490 / 55488 | 1 | 379.7 | 379.7 | 24.5 | 62.23 | 23628.731 | 33.6137 | 33.6137 | 24.5 | 62.23 | 2091.7786 | -4.6 | -11.684 | -4436.4148 | -392.7421
2 | Trim Lead tape in middle | 55298 / 55406 | 1 | 4 | 4 | 24.5 | 62.23 | 248.92 | 4.4444 | 4.4444 | 24.5 | 62.23 | 276.5778 | -4.6 | -11.684 | -46.736 | -51.9289
| wt. goal | 3240
| currently at | 3475.2 | -235.2 | grams: diff
1 | Assy., Aft Fairing | 55497
| Subtotal Weight | 5071.3 | gm | grams: diff
| Scale Weight | 4644.4 | gm | grams: -426.9
2 | Aft Fairing | 49896 | 1 | 2464 | 2464 | 54.37 | 138.0998 | 340277.9072 | 1493.5227 | 1493.5227 | 54.37 | 138.0998 | 206255.1929 | -0.02 | -0.0508 | -125.1712 | -75.871
2 | Wing, Starboard 1.5m - sent 27 Jan 09 |  | 1 | 680.5 | 680.5 | 50.942 | 129.3927 | 88051.7187 | 695 | 695 | 50.942 | 129.3927 | 89927.9126 | 0 | 0 | 0 | 0
2 | Wing, Port 1.5m - sent 27 Jan 09 |  | 1 | 637.8 | 637.8 | 50.942 | 129.3927 | 82526.6513 | 695 | 695 | 50.942 | 129.3927 | 89927.9126 | 0 | 0 | 0 | 0
2 | Wing, Starboard - 1m, internal mnt | 49838 | 0 | 433.6 | 0 | 50.942 | 129.3927 | 0 | 504.7 | 0 | 50.942 | 129.3927 | 0 | 0 | 0 | 0 | 0
2 | Wing, Port - 1m, internal mnt | 49838 | 0 | 447.1 | 0 | 50.942 | 129.3927 | 0 | 504.7 | 0 | 50.942 | 129.3927 | 0 | 0 | 0 | 0 | 0
2 | (8x) 10-32 x 1/2 PHMS PH + #10 FW + #10 LW, wings to fairing | 55241 / 55242 | 1 | 26.6 | 26.6 | 50 | 127 | 3378.2 | 3.3138 | 3.3138 | 50 | 127 | 420.8546 | 0 | 0 | 0 | 0
2 | Rudder | 49839 | 1 | 118.8 | 118.8 | 67.578 | 171.6481 | 20391.7967 | 98.1818 | 98.1818 | 67.578 | 171.6481 | 16852.7245 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 2" FHMS, rudder to rudder shoe |  | 1 | 20.3 | 20.3 | 67.75 | 172.085 | 3493.3255 | 2.529 | 2.529 | 67.75 | 172.085 | 435.1969 | 0 | 0 | 0 | 0
2 | Trim lead at rudder tube (5" strip) | 52442 / 55487 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | -1.65 | -4.191 | 0 | 0
2 | Trim Lead Tape | 55298 / 55406 | 0 | 0.5 | 0 | 62.2 | 157.988 | 0 | 0.55556 | 0 | 62.2 | 157.988 | 0 | -1.65 | -4.191 | 0 | 0
2 | Assy., Antenna | 49841
| Subtotal Weight | 535 | gm | grams: diff
| Scale Weight | 532.3 | gm | grams: -2.7
3 | Antenna Assembly GPSI (All parts) - short w/Ti pins | 52336 / 52341 / 52454 / 52476 / 55361 | 1 | 535 | 535 | 70.18 | 178.2572 | 95367.602 | 355.5 | 355.5 | 81.305 | 206.5147 | 73415.9759 | 0 | 0 | 0 | 0
3 | Cap (DGO protector during build - do not weigh) | 52481 | 1 | bar: grams
2 | Assy., Aanderaa Optode Oxygen Sensor | 55841
| Subtotal Weight | 0 | gm | grams: diff
| Scale Weight | 0 | gm | grams: 0
3 | O2 Sensor Assy. (Aanderaa optode, mount base, mount leg, SS hose clamp, (4x) 10-32 x 1/2 FHMS) | 52220 / 52485 / 52486 / 55239 / 55321 | 1 | 0 | 0 | 56.37 | 143.1798 | 0 | 122 | 122 | 56.37 | 143.1798 | 17467.9356 | 2.5 | 6.35 | 0 | 774.7
3 | O2 Sensor Cable Assembly | 55469 | 1 | 0 | 0 | 50 | 127 | 0 | 0 | 0 | 50 | 127 | 0 | 2 | 5.08 | 0 | 0
2 | Assy., SBE 43f Oxygen Sensor | 55482
| Subtotal Weight | 348.6 | gm | grams: diff
| Scale Weight | 348.1 | gm | grams: -0.5
3 | O2 Sensor Assy. (O2 sensor, Plenum, (2x) 6-32 x 5/8 SHCS - flow tube to sensor, (2x) O2 clamp, (4x) 6-32 x 1/2 FHMS - clamp to aft fairing) | 52356 / 52501 / 55214 / 52311 / 55210 | 1 | 283.9 | 283.9 | 56.37 | 143.1798 | 40648.7452 | 122 | 122 | 56.37 | 143.1798 | 17467.9356 | 2.5 | 6.35 | 1802.765 | 774.7
3 | O2 Sensor Cable Assembly | 52236 | 1 | 64.7 | 64.7 | 50 | 127 | 8216.9 | 52.6016 | 52.6016 | 50 | 127 | 6680.4065 | 2 | 5.08 | 328.676 | 267.2163
3 | Plug assy, SBE 43f plenum | 55518 | 1 | grams: 27.9
2 | Assy., WET Labs Optical | 55521
| Subtotal Weight | 239.7 | gm | grams: diff
| Scale Weight | 239.7 | gm | grams: 0
3 | Wet Labs sensor Assy. (WL BB2F-VMG sensor, (2x) sensor clamp, (2x) 8-32 x 5/8 FHMS - sensor clamp to fairing) | 52296 / 52312 / 55245 | 1 | 239.7 | 239.7 | 60.146 | 152.7708 | 36619.1703 | 159.3 | 159.3 | 60.146 | 152.7708 | 24336.3948 | -1.237 | -3.142 | -753.1326 | -500.5174
3 | Wet Labs sensor Assy. (WL Triplet sensor, (2x) Wet labs sensor clamp, (2x) 8-32 x 5/8 FHMS - sensor clamp to fairing) | 55381/ 52312 / 55245 | 1 | 0 | 0 | 60.146 | 152.7708 | 0 | 159.3 | 159.3 | 60.146 | 152.7708 | 24336.3948 | -1.237 | -3.142 | 0 | -500.5174
3 | Cover, WL sensor used in shipping, modified | 55520 | 1 | bar: grams
1 | Assy., Pupa | 55486
| Subtotal Weight | 36242.9 | gm | grams: diff
| Scale Weight | 36195 | gm | grams: -47.9
| Complete pupae volume - sum this sheet | 42758.0621 | cc
2 | Position B (Spare): IE55 dummy plug | 55309 | 1 | 7.3 | 7.3 | 51.75 | 131.445 | 959.5485 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | 0 | 0 | 0 | 0
2 | Position C (Optics Sensor / spare): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | -1.7 | -4.318 | 0 | -15.9766
2 | Position D (Oxygen sensor / spare): IE55 dummy plug | 55309 | 1 | 7.3 | 7.3 | 51.75 | 131.445 | 959.5485 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | -2.3 | -5.842 | -42.6466 | -21.6154
2 | Position E (Spare): IE55 dummy plug | 55309 | 1 | 7.3 | 7.3 | 51.75 | 131.445 | 959.5485 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | -1.7 | -4.318 | -31.5214 | -15.9766
2 | Position F (Spare): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | 1.7 | 4.318 | 0 | 15.9766
2 | O-Ring, E70-270, Endcap to Batt Hull | 52411 | 1 | 8 | 8 | 48.55 | 123.317 | 986.536 | 7.13 | 7.13 | 48.55 | 123.317 | 879.2502 | 0 | 0 | 0 | 0
2 | (4x) 6-32 x 1/2 SHCS w/ LW + (3x) 6-32 x 5/8 with zincs, endcap to battery hull | 55183 / 55217 / 55214 / 49884 | 1 | 13 | 13 | 48.64 | 123.5456 | 1606.0928 | 1.6195 | 1.6195 | 48.64 | 123.5456 | 200.0863 | 0 | 0 | 0 | 0
2 | O-Ring, E70-270, Bulkhead to batt. hull | 52411 | 1 | 8 | 8 | 28.72 | 72.9488 | 583.5904 | 7.13 | 7.13 | 28.72 | 72.9488 | 520.1249 | 0 | 0 | 0 | 0
2 | (4x / 4x) 6-32 x 1/2 SHCS w/ LW + 6-32 x 5/8 with zincs, bulkhead to battery hull | 55183 / 55217 / 55214 / 49884 | 1 | 15 | 15 | 25.75 | 65.405 | 981.075 | 1.8687 | 1.8687 | 25.75 | 65.405 | 122.2219 | 0 | 0 | 0 | 0
2 | (4x) 6-32 x 1/2 SHCS + #6 LW, rails to bulkhead | 55183 / 55217 | 1 | 5.2 | 5.2 | 23.41 | 59.4614 | 309.1993 | 0 | 0 | 23.41 | 59.4614 | 0 | 0 | 0 | 0 | 0
2 | O-Ring, E70-263, elec hull to bulkhead | 52410 | 1 | 6.3 | 6.3 | 26.13 | 66.3702 | 418.1323 | 5.76 | 5.76 | 26.13 | 66.3702 | 382.2924 | 0 | 0 | 0 | 0
2 | (8x) 6-32 x 1/2 SHCS + #6 LW, elec hull to bulkhead | 55183 / 55217 | 1 | 10.6 | 10.6 | 25.92 | 65.8368 | 697.8701 | 0 | 0 | 25.92 | 65.8368 | 0 | 0 | 0 | 0 | 0
2 | Assy., Aft Endcap | 49891
| Subtotal Weight | 7248.2 | gm | grams: diff
| Scale Weight | 7251 | gm | grams: 2.8
3 | Hydraulic oil: End Cap Without Oil | 5503 | gm
| End Cap With Oil | 6654 | gm
| Oil Added To System | 1151 | gm
| Volume of Oil In System | 1325.5787 | cc
3 | Assy., Hydraulics drive system | 55480
| Subtotal Weight | 3960.6 | gm | grams: diff
| Scale Weight | 3963.4 | gm | grams: 2.8
4 | Endcap, Aft | 49815 | 1 | 2033.2 | 2033.2 | 50.014 | 127.0356 | 258288.7006 | 2885 | 2885 | 49.752 | 126.3701 | 364577.6808 | 0 | 0 | 0 | 0
4 | (2x) Roll pin, 5/32 OD x 3/8, Pump anti rotation | 55248 | 1 | 1 | 1 | 48.41 | 122.9614 | 122.9614 | 0 | 0 | 48.41 | 122.9614 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/8 NPT x 1/4 tube, alum | 52432 | 1 | 8.3 | 8.3 | 49.3 | 125.222 | 1039.3426 | 0 | 0 | 49.3 | 125.222 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/8 NPT x 1/4 tube, SS | 52429 | 1 | 27.6 | 27.6 | 49.3 | 125.222 | 3456.1272 | 0 | 0 | 49.3 | 125.222 | 0 | 0.5 | 1.27 | 35.052 | 0
4 | Assy, Solenoid Valve, Skinner | 52254 | 1 | 612.5 | 612.5 | 49.02 | 124.5108 | 76262.865 | 0 | 0 | 49.02 | 124.5108 | 0 | 0.6 | 1.524 | 933.45 | 0
4 | Adaptor, Male, 1/8 NPT x 1/4 M, SS | 52430 | 1 | 10.6 | 10.6 | 50.43 | 128.0922 | 1357.7773 | 0 | 0 | 50.43 | 128.0922 | 0 | 0.5 | 1.27 | 13.462 | 0
4 | Branch Tee, Male, 1/4 tube x 1/8 NPT x 1/4 tube, alum | 52433 | 1 | 19 | 19 | 48.79 | 123.9266 | 2354.6054 | 0 | 0 | 48.79 | 123.9266 | 0 | -0.5 | -1.27 | -24.13 | 0
4 | Tubing, 1/4 DIA x 8" | 52419 | 1 | 3.3 | 3.3 | 47 | 119.38 | 393.954 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Tubing, 1/4 Dia x 6.0" | 52419 | 1 | 2.5 | 2.5 | 47 | 119.38 | 298.45 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Check Valve, Kepner 204 A-1 | 52448 | 1 | 22.8 | 22.8 | 49.1 | 124.714 | 2843.4792 | 0 | 0 | 49.1 | 124.714 | 0 | -0.5 | -1.27 | -28.956 | 0
4 | Union Elbow, Male, 1/4 tube x 1/4 tube, alum | 52435 | 1 | 16.3 | 16.3 | 50.43 | 128.0922 | 2087.9029 | 0 | 0 | 50.43 | 128.0922 | 0 | -0.5 | -1.27 | -20.701 | 0
4 | Tubing, 1/4 DIA x 9.625" | 52419 | 1 | 3.9 | 3.9 | 47 | 119.38 | 465.582 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Run Tee, Male,1/4 tube x 1/4 tube x 1/8 NPT, alum | 52428 | 1 | 19 | 19 | 49.3 | 125.222 | 2379.218 | 0 | 0 | 49.3 | 125.222 | 0 | 0 | 0 | 0 | 0
4 | Tubing, 1/4 Dia x 4.375" | 52419 | 1 | 1.8 | 1.8 | 47 | 119.38 | 214.884 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Seal Washer | 52350 | 1 | 0.9 | 0.9 | 50.9 | 129.286 | 116.3574 | 0 | 0 | 50.9 | 129.286 | 0 | 0 | 0 | 0 | 0
4 | LeDuc Output Screw | 52349 | 1 | 7.7 | 7.7 | 51.04 | 129.6416 | 998.2403 | 0 | 0 | 51.04 | 129.6416 | 0 | 0 | 0 | 0 | 0
4 | Hydraulic Pump, LeDuc | 52348 | 1 | 319.1 | 319.1 | 49.58 | 125.9332 | 40185.2841 | 0 | 0 | 49.58 | 125.9332 | 0 | 0 | 0 | 0 | 0
4 | Hydraulic endcap | 49803 | 1 | 110.6 | 110.6 | 47.72 | 121.2088 | 13405.6933 | 0 | 0 | 47.72 | 121.2088 | 0 | 0 | 0 | 0 | 0
4 | Ball Bearing, 8mm ID x 16mm OD | 52459 | 1 | 3.1 | 3.1 | 47.69 | 121.1326 | 375.5111 | 0 | 0 | 47.69 | 121.1326 | 0 | 0 | 0 | 0 | 0
4 | O-Ring, N70-032, hydraulic endcap | 52407 | 1 | 0.5 | 0.5 | 48.23 | 122.5042 | 61.2521 | 0 | 0 | 48.23 | 122.5042 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/16 NPT x 1/8 tube, brass | 52436 | 1 | 16.1 | 16.1 | 47.75 | 121.285 | 1952.6885 | 0 | 0 | 47.75 | 121.285 | 0 | 1.15 | 2.921 | 47.0281 | 0
4 | Tubing, 1/8 Dia. X 10.5" | 52418 | 1 | 1 | 1 | 45.17 | 114.7318 | 114.7318 | 0 | 0 | 45.17 | 114.7318 | 0 | 0 | 0 | 0 | 0
4 | Magnetic Seal | 52353 | 1 | 13.9 | 13.9 | 47.69 | 121.1326 | 1683.7431 | 0 | 0 | 47.69 | 121.1326 | 0 | 0 | 0 | 0 | 0
4 | (6x) 6-32 x 1/2 SHCS + #6 LW, hydraulic encap to encap | 55183 / 55217 | 1 | 7.9 | 7.9 | 48.12 | 122.2248 | 965.5759 | 0 | 0 | 48.12 | 122.2248 | 0 | 0 | 0 | 0 | 0
4 | Lee Plug Set | 52445 / 52446 | 1 | 0.4 | 0.4 | 51.512 | 130.8405 | 52.3362 | 0 | 0 | 51.512 | 130.8405 | 0 | -0.5 | -1.27 | -0.508 | 0
4 | Check Valve (altered item Kepner 1106 A-1-1) + O-Ring (N70 3-906) | 49852 / 52408 | 1 | 47.7 | 47.7 | 52.47 | 133.2738 | 6357.1603 | 0 | 0 | 52.47 | 133.2738 |  | 0 | 0 | 0 | 0
4 | Bladder, External | 52441 | 1 | 112.7 | 817.3705 | 54.613 | 138.717 | 113383.1976 | 924.2519 | 924.2519 | 54.613 | 138.717 | 128209.4642 | 0 | 0 | 0 | 0
4 | Nut, Bladder | 49802 | 1 | 108.1 | 108.1 | 51.6 | 131.064 | 14168.0184 | 39.8452 | 39.8452 | 51.6 | 131.064 | 5222.27 | 0 | 0 | 0 | 0
4 | Position A (Comms): IE55 w o-ring, CA W17, J1 | 52252 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | 1.7 | 4.318 | 51.816 | 3.4242
4 | Position B (Optode / Spare): port plug w/ o-ring | 52160 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | 0 | 0 | 0 | 0
4 | Position C (Optics Sensor / spare): IE55 w o-ring CA W30, J5 | 52160 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | -1.7 | -4.318 | -51.816 | -3.4242
4 | Position D (Oxygen sensor / spare): port plug w/ o-ring | 52160 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | -2.3 | -5.842 | -70.104 | -4.6327
4 | Position E (Spare): IE55 w/ o-ring | 52160 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | -1.7 | -4.318 | -51.816 | -3.4242
4 | Position F (Spare): IE55 w/ o-ring | 52160 | 1 | 12 | 12 | 51.512 | 130.8405 | 1570.0858 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | 1.7 | 4.318 | 51.816 | 3.4242
4 | Iridium/GPS Ant., conn/o-ring/ CA W5, J2 | 52241 | 1 | 50.1 | 50.1 | 51.512 | 130.8405 | 6555.108 | 6.8 | 6.8 | 51.512 | 130.8405 | 889.7153 | 0 | 0 | 0 | 0
4 | Vacuum/pressure relief valve (deep sea power) | 52449 | 1 | 11.6 | 11.6 | 51.512 | 130.8405 | 1517.7496 | 2.7 | 2.7 | 51.512 | 130.8405 | 353.2693 | 2.3 | 5.842 | 67.7672 | 15.7734
4 | Pump Plate, Inner | 49804 | 1 | 40.3 | 40.3 | 47.3 | 120.142 | 4841.7226 | 0 | 0 | 47.3 | 120.142 | 0 | 0.75 | 1.905 | 76.7715 | 0
4 | Assy, Main pump motor w/ CA W8,W18 | 52253 | 1 | 130.5 | 130.5 | 48.5 | 123.19 | 16076.295 | 0 | 0 | 48.5 | 123.19 | 0 | 1.8 | 4.572 | 596.646 | 0
4 | (3x) M2 x 6 FHMS (motor to interface plate) | 55203 | 1 | 0.4 | 0.4 | 47.3 | 120.142 | 48.0568 | 0 | 0 | 47.3 | 120.142 | 0 | 1.8 | 4.572 | 1.8288 | 0
4 | Bore Reducer | 52351 | 1 | 0.6 | 0.6 | 46.77 | 118.7958 | 71.2775 | 0 | 0 | 46.77 | 118.7958 | 0 | 0 | 0 | 0 | 0
4 | Pulley, LeDuc pump motor | 52352 | 1 | 5.7 | 5.7 | 46.77 | 118.7958 | 677.1361 | 0 | 0 | 46.77 | 118.7958 | 0 | 1.85 | 4.699 | 26.7843 | 0
4 | (6x) 4-40 x 3/8 FHMS, hyd end cap to inner plate | 55229 | 1 | 2.9 | 2.9 | 47.28 | 120.0912 | 348.2645 | 0 | 0 | 47.28 | 120.0912 | 0 | 0 | 0 | 0 | 0
4 | Pulley, LeDuc pump | 49806 | 1 | 91.5 | 91.5 | 46.77 | 118.7958 | 10869.8157 | 0 | 0 | 46.77 | 118.7958 | 0 | 0 | 0 | 0 | 0
4 | Timing Belt | 52357 | 1 | 3.5 | 3.5 | 46.77 | 118.7958 | 415.7853 | 0 | 0 | 46.77 | 118.7958 | 0 | 0.5 | 1.27 | 4.445 | 0
4 | Pump Plate, Outer | 49805 | 1 | 18 | 18 | 46.5 | 118.11 | 2125.98 | 0 | 0 | 46.5 | 118.11 | 0 | 0.75 | 1.905 | 34.29 | 0
4 | (4x) 4-40 x 3/16 SHCS + #4 LW, outer plate to inner plate | 55296 | 1 | 2 | 2 | 46.625 | 118.4275 | 236.855 | 0 | 0 | 46.625 | 118.4275 | 0 | 0 | 0 | 0 | 0
3 | Assy., Boost Pump | 52378
| Subtotal Weight | 382.6 | gm | grams: diff
| Scale Weight | 382.8 | gm | grams: 0.2
4 | Pump Bracket | 49808 | 1 | 25.1 | 25.1 | 48.91 | 124.2314 | 3118.2081 | 0 | 0 | 48.91 | 124.2314 | 0 | 0 | 0 | 0 | 0
4 | Motor, Chrge Pump w/ W34 | 52288 | 1 | 60.3 | 60.3 | 48.9 | 124.206 | 7489.6218 | 0 | 0 | 48.9 | 124.206 | 0 | 1 | 2.54 | 153.162 | 0
4 | (2x) M2 x 6 FHMS, motor to bracket | 55203 | 1 | 0.2 | 0.2 | 48.9 | 124.206 | 24.8412 | 0 | 0 | 48.9 | 124.206 | 0 | 0 | 0 | 0 | 0
4 | Magnet Rotor Housing | 52457 | 1 | 10.3 | 10.3 | 48.9 | 124.206 | 1279.3218 | 0 | 0 | 48.9 | 124.206 | 0 | -0.4 | -1.016 | -10.4648 | 0
4 | (4x) 6-32 x 1/4 PHMS, bracket to housing | 55209 | 1 | 3 | 3 | 48.9 | 124.206 | 372.618 | 0 | 0 | 48.9 | 124.206 | 0 | 0 | 0 | 0 | 0
4 | Rotary Magnet | 52455 | 1 | 17.8 | 17.8 | 48.9 | 124.206 | 2210.8668 | 0 | 0 | 48.9 | 124.206 | 0 | -0.4 | -1.016 | -18.0848 | 0
4 | Pumphead, Micropump + mounting plate + (3x) 4-40 x 3/8 FHMS | 52360 / 52456 / 55229 | 1 | 224.5 | 224.5 | 48.9 | 124.206 | 27884.247 | 0 | 0 | 48.9 | 124.206 | 0 | -0.8 | -2.032 | -456.184 | 0
4 | (4x) 4-40 x 3/16 SHCS + #4 LW, housing to flange | 55296 / 55233 | 1 | 2 | 2 | 48.9 | 124.206 | 248.412 | 0 | 0 | 48.9 | 124.206 | 0 | -0.5 | -1.27 | -2.54 | 0
4 | (2x) Elbow, 1/8 npt x 1/4 tube, alum | 52431 | 1 | 25.1 | 25.1 | 48.9 | 124.206 | 3117.5706 | 0 | 0 | 48.9 | 124.206 | 0 | -0.9 | -2.286 | -57.3786 | 0
4 | Filter, Balston | 52425 | 1 | 14.3 | 14.3 | 47 | 119.38 | 1707.134 | 0 | 0 | 47 | 119.38 |  | 0 | 0 | 0 | 0
3 | Assy., Hydraulic Reservoir | 52346
| Subtotal Weight | 1158.4 | gm | grams: diff
| Scale Weight | 1155.2 | gm | grams: -3.2
4 | Piston, Diaphragm | 49811 | 1 | 139.1 | 139.1 | 45 | 114.3 | 15899.13 | 0 | 0 | 45 | 114.3 | 0 | 0 | 0 | 0 | 0
4 | Diaphragm, Internal (Bellofram) | 52359 | 1 | 95.9 | 542.2295 | 45.5531 | 115.7048 | 62738.5626 | 0 | 0 | 45.5531 | 115.7048 | 0 | 0 | 0 | 0 | 0
4 | Sikaflex Adhesive Compound | 55459 | 1 | 26.7 | 26.7 | 45 | 114.3 | 3051.81 | 0 | 0 | 45 | 114.3 | 0 | 0 | 0 | 0 | 0
4 | Cylinder Head, Diaphragm | 49813 | 1 | 321.6 | 321.6 | 43.9 | 111.506 | 35860.3296 | 0 | 0 | 43.9 | 111.506 | 0 | 0 | 0 | 0 | 0
4 | Branch Tee, Male, 1/4 tube x 1/8 NPT x 1/4 tube, brass + Plastic plug (X59P4) | 52443 | 1 | 23.5 | 23.5 | 43.24 | 109.8296 | 2580.9956 | 0 | 0 | 43.24 | 109.8296 | 0 | -2.39 | -6.0706 | -142.6591 | 0
4 | Elbow, 1/8 NPT M x 1/8 NPT M, brass | 52426 | 1 | 26.9 | 26.9 | 43.24 | 109.8296 | 2954.4162 | 0 | 0 | 43.24 | 109.8296 | 0 | 2.39 | 6.0706 | 163.2991 | 0
4 | Kepner Chk Valve 404A-1-10 | 52447 | 1 | 25.6 | 25.6 | 43.24 | 109.8296 | 2811.6378 | 0 | 0 | 43.24 | 109.8296 | 0 | 0.8 | 2.032 | 52.0192 | 0
4 | Elbow, 1/8 npt x 1/8 tube, alum | 52427 | 1 | 11 | 11 | 43.24 | 109.8296 | 1208.1256 | 0 | 0 | 43.24 | 109.8296 | 0 | -0.1 | -0.254 | -2.794 | 0
4 | Bleed screw, 10-32 x .125 PHMS with seal | 55477 | 1 | 1.7 | 1.7 | 43.24 | 109.8296 | 186.7103 | 0 | 0 | 43.24 | 109.8296 | 0 | 3 | 7.62 | 12.954 | 0
4 | Cylinder, Diaphragm | 49812 | 1 | 425.8 | 425.8 | 45.3 | 115.062 | 48993.3996 | 0 | 0 | 45.3 | 115.062 | 0 | 0 | 0 | 0 | 0
4 | (8x) 5-40 x 3/8 FHMS, head to cylinder | 55251 | 1 | 7.4 | 7.4 | 43.56 | 110.6424 | 818.7538 | 0 | 0 | 43.56 | 110.6424 | 0 | 0 | 0 | 0 | 0
4 | Standoff (starboard) | 55299 | 1 | 12.6 | 12.6 | 47.75 | 121.285 | 1528.191 | 0 | 0 | 47.75 | 121.285 | 0 | 2.6 | 6.604 | 83.2104 | 0
4 | Standoff (port) | 55299 | 1 | 12.6 | 12.6 | 47.75 | 121.285 | 1528.191 | 0 | 0 | 47.75 | 121.285 | 0 | -2.6 | -6.604 | -83.2104 | 0
4 | (4x) 6-32 x 3/8 SHCS, standoff to cylinder | 55183 | 1 | 4.1 | 4.1 | 47.5 | 120.65 | 494.665 | 0 | 0 | 47.5 | 120.65 | 0 | 0 | 0 | 0 | 0
4 | (2x) Saddles, mounting, Aft PCB + (4x) Nylon 4-40 x 1/4 PHMS | 52373 / 55257 | 1 | 5 | 5 | 45.5 | 115.57 | 577.85 | 0 | 0 | 45.5 | 115.57 | 0 | 3.2 | 8.128 | 40.64 | 0
4 | (2x) Saddles, mounting, C PCB + (4x) Nylon 4-40 x 1/4 PHMS | 49888 / 55257 | 1 | 6.3 | 6.3 | 45.5 | 115.57 | 728.091 | 0 | 0 | 45.5 | 115.57 | 0 | 1.5 | 3.81 | 24.003 | 0
4 | (2x) Saddles, mounting, T PCB + (4x) Nylon 4-40 x 1/4 PHMS | 49888 / 55257 | 1 | 6.3 | 6.3 | 45.5 | 115.57 | 728.091 | 0 | 0 | 45.5 | 115.57 | 0 | -1.5 | -3.81 | -24.003 | 0
4 | (2x) Saddles, mounting, O2 PCB + (4x) Nylon 4-40 x 1/4 PHMS (only used with SBE 43f) | 49888 / 55257 | 1 | 6.3 | 6.3 | 45.5 | 115.57 | 728.091 | 0 | 0 | 45.5 | 115.57 | 0 | -3.2 | -8.128 | -51.2064 | 0
Weigh Hydraulic res. Assy, then Oil fill & add sub-assy fasteners
3 | Linear Poteniometer A w/ cable assy (W21), Linear pot clamp, 3-48 x 1/2 SHCS, Spring, Linear pot rod end | 52255 / 49814 / 55220 | 1 | 39.4 | 39.4 | 47.75 | 121.285 | 4778.629 | 0 | 0 | 47.75 | 121.285 | 0 | 2.5 | 6.35 | 250.19 | 0
3 | Linear Poteniometer B w/ cable assy (W20), Linear pot clamp, 3-48 x 1/2 SHCS, Spring, Linear pot rod end | 52256 / 49814 / 55220 | 1 | 37.6 | 37.6 | 47.75 | 121.285 | 4560.316 | 0 | 0 | 47.75 | 121.285 | 0 | 2.5 | 6.35 | 238.76 | 0
3 | (2x) 5-40 x 1/2 SHCS + #5 LW, Clamp to cylinder | 55250 / 55212 / 55225 | 1 | 2.2 | 2.2 | 47.64 | 121.0056 | 266.2123 | 0 | 0 | 47.64 | 121.0056 | 0 | 0 | 0 | 0 | 0
3 | (2x) 6-32 x 1/2 SHCS + #6 LW, boost pump bracket to endcap | 55183 / 55217 | 1 | 2.7 | 2.7 | 48.28 | 122.6312 | 331.1042 | 0 | 0 | 48.28 | 122.6312 | 0 | 0 | 0 | 0 | 0
3 | (4x) 6-32 x 1/2 SHCS + #6 LW, cylinder standoff to endcap | 55183 / 55217 | 1 | 5.4 | 5.4 | 47.9 | 121.666 | 656.9964 | 0 | 0 | 47.9 | 121.666 | 0 | 0 | 0 | 0 | 0
aft endcap electronic parts - mount with CT sail
3 | Aft Terminal PCB | 55302 | 1 | 24.8 | 24.8 | 45.5 | 115.57 | 2866.136 | 0 | 0 | 45.5 | 115.57 | 0 | 3.4 | 8.636 | 214.1728 | 0
3 | SBE Conductivity PCB (included with sensors) | 55385 | 1 | 33.7 | 33.7 | 45.5 | 115.57 | 3894.709 | 0 | 0 | 45.5 | 115.57 | 0 | 1.7 | 4.318 | 145.5166 | 0
3 | Conductivity, PCB jumper, CA W26 | 52261 | 1 | 4 | 4 | 44.4 | 112.776 | 451.104 | 0 | 0 | 44.4 | 112.776 | 0 | 1.5 | 3.81 | 15.24 | 0
3 | Shield, Radiation, C PCB + (4x) 2-56 x 1/4 BHCS | 49887 / 55221 | 1 | 21.3 | 21.3 | 45.5 | 115.57 | 2461.641 | 0 | 0 | 45.5 | 115.57 | 0 | 1.9 | 4.826 | 102.7938 | 0
3 | SBE Temperature PCB (included with sensors) | 55388 | 1 | 22.2 | 22.2 | 45.5 | 115.57 | 2565.654 | 0 | 0 | 45.5 | 115.57 | 0 | -1.7 | -4.318 | -95.8596 | 0
3 | Temperature, PCB jumper, CA W25 | 52260 | 1 | 5.1 | 5.1 | 44.4 | 112.776 | 575.1576 | 0 | 0 | 44.4 | 112.776 | 0 | 0.5 | 1.27 | 6.477 | 0
3 | Shield, Radiation, T PCB + (4x) 2-56 x 1/4 BHCS | 52371 / 55221 | 1 | 17.2 | 17.2 | 45.5 | 115.57 | 1987.804 | 0 | 0 | 45.5 | 115.57 | 0 | -1.9 | -4.826 | -83.0072 | 0
3 | SBE Oxygen PCB (only used with SBE 43f) ((included with sensors) | 55524 | 1 | 23.1 | 23.1 | 45.5 | 115.57 | 2669.667 | 0 | 0 | 45.5 | 115.57 | 0 | -3.4 | -8.636 | -199.4916 | 0
3 | Oxygen, PCB jumper, CA W33 (only used with SBE 43f) | 52265 | 1 | 5.4 | 5.4 | 44.4 | 112.776 | 608.9904 | 0 | 0 | 44.4 | 112.776 | 0 | 0 | 0 | 0 | 0
3 | Ribbon Cable (W13) (gray) | 52248 | 1 | 28.4 | 28.4 | 38.625 | 98.1075 | 2786.253 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.5 | 8.89 | 252.476 | 0
3 | Ribbon Cable (W14) (rainbow) | 52249 | 1 | 37 | 37 | 38.625 | 98.1075 | 3629.9775 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.5 | 8.89 | 328.93 | 0
3 | CT Sensor Sail Assembly (stanchion, guard, nut, nacelle, (6x) 6-32 x 5/16 FHMS, o-ring (E70-016), o-ring (E70-012), SBE components) | 52368 | 1 | 286.1 | 286.1 | 47.864 | 121.5746 | 34782.4816 | 184.027 | 184.027 | 47.864 | 121.5746 | 22373.0016 | 7.058 | 17.9273 | 5129.0063 | 3299.1109
2 | Assy., Mass shifter | 49889
| Subtotal Weight | 12455.7 | gm | grams: diff
| Scale Weight | 12464 | gm | grams: 8.3
3 | Assy., Bulkhead | 55843
4 | Hull Bulkhead | 49819 | 1 | 1301.7 | 1301.7 | 27.322 | 69.3979 | 90335.2204 | 2669.8 | 2669.8 | 27.506 | 69.8652 | 186526.2178 | 0 | 0 | 0 | 0
pressure test with partial fwd hull 49877 and complete 54844/5
3 | Centering Spring, Roll Bearing | 49845 | 1 | 7.5 | 7.5 | 27.58 | 70.0532 | 525.399 | 0 | 0 | 27.58 | 70.0532 | 0 | 0 | 0 | 0 | 0
3 | Wire Bridge with vinyl tape wrap | 49894 | 1 | 8.6 | 8.6 | 26.8 | 68.072 | 585.4192 | 0 | 0 | 26.8 | 68.072 | 0 | 0 | 0 | 0 | 0
3 | (2x) 6-32 x 5/16 FHMS, wire bridge to bulkhead | 55208 | 1 | 1.1 | 1.1 | 26.8 | 68.072 | 74.8792 | 0 | 0 | 26.8 | 68.072 | 0 | 0 | 0 | 0 | 0
3 | Snap Ring, Roll Bearing | 52402 | 1 | 53.2 | 53.2 | 27.8 | 70.612 | 3756.5584 | 0 | 0 | 27.8 | 70.612 | 0 | 0 | 0 | 0 | 0
3 | (4x) 6-32 x 1/2 FHMS, roll gear block to bulkhead | 55210 | 1 | 3.7 | 3.7 | 26.28 | 66.7512 | 246.9794 | 0 | 0 | 26.28 | 66.7512 | 0 | -3.14 | -7.9756 | -29.5097 | 0
3 | Spur Gear, Delrin, Roll Pot. | 52468 | 1 | 6.9 | 6.9 | 27.25 | 69.215 | 477.5835 | 0 | 0 | 27.25 | 69.215 | 0 | -2.43 | -6.1722 | -42.5882 | 0
3 | Battery Pack, HV with CA W2 + shim | 52238 | 1 | 6487.9 | 6487.9 | 36.8465 | 93.5902 | 607203.6699 | 0 | 0 | 36.8465 | 93.5902 | 0 | -0.333 | -0.84582 | -5487.5956 | 0
3 | Bottom Ballast Weight | 52496 | 1 | 1164.9 | 1164.9 | 36.8465 | 93.5902 | 109023.1901 | 0 | 0 | 36.8465 | 93.5902 | 0 | -3.75 | -9.525 | -11095.6725 | 0
3 | (8x) 4-40 x 3/8 SHCS, ballast to batt encl endplates | 55228 | 1 | 4.8 | 4.8 | 36.8465 | 93.5902 | 449.2328 | 0 | 0 | 36.8465 | 93.5902 | 0 | -3.75 | -9.525 | -45.72 | 0
4 | Assy., Paine Pressure | 55484
| Subtotal Weight | 128.4 | gm | grams: diff
| Scale Weight | 128.4 | gm | grams: 0
5 | Paine Sensor w/ CA W12 | 52247 | 1 | 92.3 | 92.3 | 26.84 | 68.1736 | 6292.4233 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -117.221 | 0
5 | Male Connector, 1/16 NPT x 1/8 tube | 52438 | 1 | 15.3 | 15.3 | 26.84 | 68.1736 | 1043.0561 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -19.431 | 0
5 | Female Connector, 7/16-20 UNF x 1/8 tube + O-Ring (E70-012) | 49816 / 52409 | 1 | 18.6 | 18.6 | 26.84 | 68.1736 | 1268.029 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -23.622 | 0
5 | Tubing, 0.125 D x 1.1 L SS | 52418 | 1 | 1.9 | 1.9 | 26.84 | 68.1736 | 129.5298 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -2.413 | 0
5 | FS-1265 Fluid in Pressure S. | 55472 | 1 | 0.3 | 0.3 | 26.84 | 68.1736 | 20.4521 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -0.381 | 0
4 | Assy., Druck Pressure Sensor | 55485
| Subtotal Weight | 0 | gm | grams: diff
| Scale Weight | 0 | gm | grams: 0
5 | Druck Sensor w/ CA W12 | 55517 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Male Connector, 1/16 NPT x 1/8 tube | 52438 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Male connector, 1/4 NPT x 1/8 tube | 52347 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Tubing, 0.125 D x 1.2 L SS | 52418 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | FS-1265 Fluid in Pressure S. | 55472 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
3 | Assy., Pitch Mechanism | 55531
| Subtotal Weight | 2937.4 | gm | grams: diff
| Scale Weight | 2938.2 | gm | grams: 0.8
4 | (2x) Bracket, Linear Brng Rail-fwd | 49822 | 1 | 30.6 | 30.6 | 28.58 | 72.5932 | 2221.3519 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -169.4383 | 0
4 | (2x) Bracket, Linear Brng Rail-aft | 49822 | 1 | 30.7 | 30.7 | 41.72 | 105.9688 | 3253.2422 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -169.992 | 0
4 | (4x) Linear Bearing Car | 55532 | 1 | 123 | 123 | 36.8465 | 93.5902 | 11511.591 | 0 | 0 | 36.8465 | 93.5902 | 0 | -1.75 | -4.445 | -546.735 | 0
4 | (2x) Linear Brng Rail; THK RSR12VM | 52440 | 1 | 431.5 | 431.5 | 35.2 | 89.408 | 38579.552 | 0 | 0 | 35.2 | 89.408 | 0 | -2.18 | -5.5372 | -2389.3018 | 0
4 | (4x) 6-32 x 5/8 FHMS + 6-32 nylok, linear rail to bracket - fwd | 55211 / 55216 | 1 | 8.6 | 8.6 | 28.58 | 72.5932 | 624.3015 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -47.6199 | 0
4 | (4x) 6-32 x 5/8 FHMS + 6-32 nylok, linear rail to bracket - aft | 55211 / 55216 | 1 | 8.6 | 8.6 | 41.72 | 105.9688 | 911.3317 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -47.6199 | 0
4 | (2x) Batt. encl side plate | 52492 | 1 | 168.8 | 168.8 | 36.8465 | 93.5902 | 15798.0209 | 0 | 0 | 36.8465 | 93.5902 | 0 | -2.14 | -5.4356 | -917.5293 | 0
4 | (16x) M6 x 6 FHCS, batt. encl side plate to bearing car | 55202 | 1 | 5.9 | 5.9 | 36.8465 | 93.5902 | 552.182 | 0 | 0 | 36.8465 | 93.5902 | 0 | -1.75 | -4.445 | -26.2255 | 0
4 | Batt. encl endplate, fwd | 49869 | 1 | 111 | 111 | 32.9085 | 83.5877 | 9278.2293 | 0 | 0 | 32.9085 | 83.5877 | 0 | -0.33 | -0.8382 | -93.0402 | 0
4 | Batt. encl endplate, aft | 49869 | 1 | 110.4 | 110.4 | 40.7845 | 103.5927 | 11436.6331 | 0 | 0 | 40.7845 | 103.5927 | 0 | -0.33 | -0.8382 | -92.5373 | 0
4 | (8x) 4-40 x 3/8 FHCS, side plate to batt encl ends | 55229 / 55280 | 1 | 4 | 4 | 36.8465 | 93.5902 | 374.3607 | 0 | 0 | 36.8465 | 93.5902 | 0 | -2.14 | -5.4356 | -21.7424 | 0
4 | Fwd Endplate, MassShifter | 49824 | 1 | 280.1 | 280.1 | 27.63 | 70.1802 | 19657.474 | 0 | 0 | 27.63 | 70.1802 | 0 | 0 | 0 | 0 | 0
4 | Bearing, Ball Screw Shaft-fwd, .375 ID x .875 OD | 52458 | 1 | 11.5 | 11.5 | 27.75 | 70.485 | 810.5775 | 0 | 0 | 27.75 | 70.485 | 0 | 3.25 | 8.255 | 94.9325 | 0
4 | (2x) Retaining Ring, internal spiral, Bearing-fwd | 55271 | 1 | 1.2 | 1.2 | 27.75 | 70.485 | 84.582 | 0 | 0 | 27.75 | 70.485 | 0 | 3.25 | 8.255 | 9.906 | 0
4 | Aft End Plate, MassShifter | 49821 | 1 | 96.3 | 96.3 | 42.9 | 108.966 | 10493.4258 | 0 | 0 | 42.9 | 108.966 | 0 | 0 | 0 | 0 | 0
4 | Bearing, Ball Screw Shaft-aft, .375 ID x .875 OD | 52458 | 1 | 11.5 | 11.5 | 42.72 | 108.5088 | 1247.8512 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 94.9325 | 0
4 | (2x) Retaining Ring, internal spiral, Bearing-aft | 55271 | 1 | 1.2 | 1.2 | 42.72 | 108.5088 | 130.2106 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 9.906 | 0
4 | Internal Ring Gear, Roll | 49854 / 52488 | 1 | 86.5 | 86.5 | 27.72 | 70.4088 | 6090.3612 | 0 | 0 | 27.72 | 70.4088 | 0 | -2.6 | -6.604 | -571.246 | 0
4 | (6x) 4-40 x 3/8 FHMS, internal gear to fwd endplate | 55229 | 1 | 3 | 3 | 27.7 | 70.358 | 211.074 | 0 | 0 | 27.7 | 70.358 | 0 | -2.6 | -6.604 | -19.812 | 0
4 | Roll Bearing; KU075 XPO | 52487 | 1 | 405.2 | 405.2 | 27.57 | 70.0278 | 28375.2646 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | Ball Screw Shaft + spacers | 49820 / 55188 | 1 | 329.7 | 329.7 | 35.2 | 89.408 | 29477.8176 | 0 | 0 | 35.2 | 89.408 | 0 | 3.2 | 8.128 | 2679.8016 | 0
4 | Worm Gear, Pitch, Brass G1035 | 49853 | 1 | 46 | 46 | 27.15 | 68.961 | 3172.206 | 0 | 0 | 27.15 | 68.961 | 0 | 3.23 | 8.2042 | 377.3932 | 0
4 | Roll Pin, pitch worm gear to shaft, 3/32 OD x 5/8 L | 55272 | 1 | 0.4 | 0.4 | 27.34 | 69.4436 | 27.7774 | 0 | 0 | 27.34 | 69.4436 | 0 | 3.225 | 8.1915 | 3.2766 | 0
4 | (3x) Retaining Plate; Roll Bearing | 49823 | 1 | 12 | 12 | 27.57 | 70.0278 | 840.3336 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | (6x) 6-32 x 1/2 FHMS, ring plate to fwd plate |  | 1 | 5.6 | 5.6 | 27.57 | 70.0278 | 392.1557 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | Ball Screw Nut THK BTP0502 | 52403 | 1 | 221.4 | 221.4 | 36.8465 | 93.5902 | 20720.8638 | 0 | 0 | 36.8465 | 93.5902 | 0 | 3.2 | 8.128 | 1799.5392 | 0
4 | Mount, Ball Nut Screw | 52375 | 1 | 14.8 | 14.8 | 36.8465 | 93.5902 | 1385.1345 | 0 | 0 | 36.8465 | 93.5902 | 0 | 2.9 | 7.366 | 109.0168 | 0
4 | (4x) 8-32 x 5/8 SHCS + #8 LW, bracket to batt end plate, fwd | 55244 / 55247 | 1 | 9.4 | 9.4 | 28.58 | 72.5932 | 682.3761 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -52.0497 | 0
4 | (4x) 8-32 x 5/8 SHCS + #8 LW, bracket to batt end plate, aft | 55244 / 55247 | 1 | 9.4 | 9.4 | 41.72 | 105.9688 | 996.1067 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -52.0497 | 0
4 | Retaining Ring, external .338 ID, Shaft-aft | 55263 | 1 | 0.1 | 0.1 | 42.72 | 108.5088 | 10.8509 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 0.8255 | 0
4 | Top Plate | 52493 | 1 | 166.9 | 166.9 | 36.8465 | 93.5902 | 15620.1995 | 0 | 0 | 36.8465 | 93.5902 | 0 | 2.78 | 7.0612 | 1178.5143 | 0
4 | (4x) 4-40 x 3/8 SHCS + #4 FW, top plate assy to batt encl end | 55228 / 55280 | 1 | 3.6 | 3.6 | 36.8465 | 93.5902 | 336.9246 | 0 | 0 | 36.8465 | 93.5902 | 0 | -2.14 | -5.4356 | -19.5682 | 0
4 | (2x) 6-32 x 3/8 SHCS + #6 FW + #6 LW, nut retainer to top plate assy | 55213 / 55217 | 1 | 2.9 | 2.9 | 36.8465 | 93.5902 | 271.4115 | 0 | 0 | 36.8465 | 93.5902 | 0 | 2.9 | 7.366 | 21.3614 | 0
4 | Motor, Pitch with C.A. W41 | 52286 | 1 | 71.3 | 71.3 | 27.06 | 68.7324 | 4900.6201 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 425.5897 | 0
4 | Interface Plate, Pitch Motor | 49850 | 1 | 1.4 | 1.4 | 27.06 | 68.7324 | 96.2254 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 8.3566 | 0
4 | (2x) M2 x 6 CHMS + M2 LW, motor to interface plate | 55201 / 55200 | 1 | 0.4 | 0.4 | 27.06 | 68.7324 | 27.493 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 2.3876 | 0
4 | Gear Block, Pitch Motor | 49825 | 1 | 23 | 23 | 27.54 | 69.9516 | 1608.8868 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 139.6238 | 0
4 | (2x) Bearings, Pitch Motor, .25 ID x .625 OD flanged | 52452 | 1 | 9.2 | 9.2 | 27.06 | 68.7324 | 632.3381 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 54.9148 | 0
4 | Spacers, Pitch motor shaft | 55187 / 55267 / 55268 / 55260 | 1 | 2.3 | 2.3 | 27.06 | 68.7324 | 158.0845 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 13.7287 | 0
4 | Worm, Pitch | 52467 | 1 | 15.9 | 15.9 | 27.06 | 68.7324 | 1092.8452 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 94.9071 | 0
4 | Roll pin, pitch worm to motor shaft, .094 OD x .375 L | 55276 | 1 | 0.2 | 0.2 | 27.06 | 68.7324 | 13.7465 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 1.1938 | 0
4 | (4x) 2-56 x 5/16 PHMS, interface plate to gear block | 55219 | 1 | 1 | 1 | 27.06 | 68.7324 | 68.7324 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 5.969 | 0
4 | Gear Block shim plates, pitch motor | 55533 | 1 | 0 | 0 | 27.54 | 69.9516 | 0 | 0 | 0 | 27.54 | 69.9516 |  | 2.39 | 6.0706 | 0 | 0
4 | (4x) Shoulder screw, .125 dia x 7/16 L | 55222 | 1 | 5.9 | 5.9 | 27.54 | 69.9516 | 412.7144 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 35.8165 | 0
4 | (4x) # 4 nylok | 55232 | 1 | 2.2 | 2.2 | 27.54 | 69.9516 | 153.8935 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 13.3553 | 0
4 | 10 turn Potentiometer w/ nut and cable W41, pitch | 52287 | 1 | 30.6 | 30.6 | 27.72 | 70.4088 | 2154.5093 | 0 | 0 | 27.72 | 70.4088 | 0 | 1.99 | 5.0546 | 154.6708 | 0
4 | Spur gear, Pitch shaft, 48 tooth, delrin | 52460 | 1 | 9 | 9 | 26.8 | 68.072 | 612.648 | 0 | 0 | 26.8 | 68.072 | 0 | 3.24 | 8.2296 | 74.0664 | 0
4 | Pitch pot gear, 72 tooth, delrin | 52461 | 1 | 13.2 | 13.2 | 26.78 | 68.0212 | 897.8798 | 0 | 0 | 26.78 | 68.0212 | 0 | 2.01 | 5.1054 | 67.3913 | 0
*
3 | Assy., Roll Motor | 55530
| Subtotal Weight | 349.6 | gm | grams: diff
| Scale Weight | 349.6 | gm | grams: 0
4 | Gear Block, Roll Motor | 49827 | 1 | 105.5 | 105.5 | 26.9 | 68.326 | 7208.393 | 0 | 0 | 26.9 | 68.326 | 0 | -2.43 | -6.1722 | -651.1671 | 0
4 | (4x) Bearing, Roll mechanism, .25 ID x .625 OD flanged | 52452 | 1 | 18.5 | 18.5 | 26.8 | 68.072 | 1259.332 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -70.9549 | 0
4 | Motor, Roll w/ shaft and C.A. W15 | 52250 | 1 | 71.4 | 71.4 | 26.8 | 68.072 | 4860.3408 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -273.8476 | 0
4 | Interface Plate, Roll Motor | 49850 | 1 | 1.4 | 1.4 | 26.8 | 68.072 | 95.3008 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -5.3696 | 0
4 | (2x) M2 x 6 CHMS + M2 LW, motor to interface plate | 55201 / 55200 | 1 | 0.4 | 0.4 | 26.8 | 68.072 | 27.2288 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -1.5342 | 0
4 | Spacers, Roll Motor Shaft | 55266 / 55267 / 55260 | 1 | 1.9 | 1.9 | 26.8 | 68.072 | 129.3368 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -7.2873 | 0
4 | Worm, Roll | 52489 | 1 | 9.6 | 9.6 | 26.8 | 68.072 | 653.4912 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -36.8198 | 0
4 | (4x) 2-56 x 5/16 PHMS, interface plate to gear block | 55219 | 1 | 1 | 1 | 26.8 | 68.072 | 68.072 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -3.8354 | 0
4 | Roll gear shaft | 49828 | 1 | 9.6 | 9.6 | 26.92 | 68.3768 | 656.4173 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -63.8861 | 0
4 | Worm Gear, Roll | 49851 | 1 | 88 | 88 | 26.86 | 68.2244 | 6003.7472 | 0 | 0 | 26.86 | 68.2244 | 0 | -2.62 | -6.6548 | -585.6224 | 0
4 | Roll Pin, Roll Gear Shaft, .063 OD x .75 L | 55275 | 1 | 0.2 | 0.2 | 26.92 | 68.3768 | 13.6754 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -1.331 | 0
4 | Spacers, Ret. Rings, roll gear shaft | 55266 / 55267 / 55268 / 55266 | 1 | 1.1 | 1.1 | 26.92 | 68.3768 | 75.2145 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -7.3203 | 0
4 | Spur Gear, Roll | 49872 | 1 | 10.4 | 10.4 | 27.64 | 70.2056 | 730.1382 | 0 | 0 | 27.64 | 70.2056 | 0 | -2.66 | -6.7564 | -70.2666 | 0
4 | Roll Pin, Roll Spur Gear, .063 dia x .500 | 55274 | 1 | 0.1 | 0.1 | 27.64 | 70.2056 | 7.0206 | 0 | 0 | 27.64 | 70.2056 | 0 | -2.66 | -6.7564 | -0.67564 | 0
4 | Potentiometer, Roll w/ nut and CA W16 | 52251 | 1 | 30.5 | 30.5 | 27.25 | 69.215 | 2111.0575 | 0 | 0 | 27.25 | 69.215 | 0 | -2.43 | -6.1722 | -188.2521 | 0
2 | Assy., Electronics | 52347
| Subtotal Weight | 3201.4 | gm | grams: diff
| Scale Weight | 3222.7 | gm | grams: 21.3
3 | Electronics rail, port | 49846 | 1 | 64.3 | 64.3 | 23.507 | 59.7078 | 3839.2103 | 0 | 0 | 23.507 | 59.7078 | 0 | 0 | 0 | 0 | 0
3 | Electronics rail, starboard | 52379 | 1 | 64.4 | 64.4 | 23.507 | 59.7078 | 3845.181 | 0 | 0 | 23.507 | 59.7078 | 0 | 0 | 0 | 0 | 0
3 | Seaglider Mainboard (includes all motherboard electronics, TT8, CF8 w/ flash disk, all permanent pigtails) | 52290-B.4 | 1 | 267.1 | 267.1 | 21.5 | 54.61 | 14586.331 | 0 | 0 | 21.5 | 54.61 | 0 | 0.7 | 1.778 | 474.9038 | 0
3 | Insulation Sheet-OEM/Main | 52370 | 1 | 3.8 | 3.8 | 21.73 | 55.1942 | 209.738 | 0 | 0 | 21.73 | 55.1942 | 0 | 0.44 | 1.1176 | 4.2469 | 0
3 | (6x) 6-32 x 1/4 PHMS, main board to rails | 55237 | 1 | 4.5 | 4.5 | 21.5 | 54.61 | 245.745 | 0 | 0 | 21.5 | 54.61 | 0 | 0.5 | 1.27 | 5.715 | 0
3 | Assy., OEM Navigation Board (includes transponder PCB, transformer/inductor, capacitors, GPS reciever, W6 GPS serial interface, DC-DC converter, Cable Assembly W7, all soldered on wiring) | 52257 | 1 | 554.8 | 554.8 | 21.5 | 54.61 | 30297.628 | 0 | 0 | 21.5 | 54.61 | 0 | -0.43 | -1.0922 | -605.9526 | 0
3 | (6x) 6-32 x 1/4 PHMS, OEM board to rails | 55237 | 1 | 4.6 | 4.6 | 21.48 | 54.5592 | 250.9723 | 0 | 0 | 21.48 | 54.5592 | 0 | -0.43 | -1.0922 | -5.0241 | 0
3 | Assy., Iridium Modem and bracket (includes satellite reciever - iridium LBT, cable assy W3, all clamps for wires, RF switch, DC block, and mounting hardware) | 52324 | 1 | 282.4 | 282.4 | 21.6 | 54.864 | 15493.5936 | 0 | 0 | 21.6 | 54.864 | 0 | 1.9 | 4.826 | 1362.8624 | 0
3 | (4x) 6-32 x 1/4 PHMS + #6 LW, modem cage to rails | 55209 / 55217 | 1 | 3.5 | 3.5 | 20.77 | 52.7558 | 184.6453 | 0 | 0 | 20.77 | 52.7558 | 0 | 0.29 | 0.7366 | 2.5781 | 0
3 | Battery Pack, LV w/ CA W1 | 52237 | 1 | 1872.5 | 1872.5 | 21.6 | 54.864 | 102732.84 | 0 | 0 | 21.6 | 54.864 | 0 | -1.81 | -4.5974 | -8608.6315 | 0
3 | Battery Pack Cage, LV | 49847 | 1 | 74.2 | 74.2 | 21.6 | 54.864 | 4070.9088 | 0 | 0 | 21.6 | 54.864 | 0 | -1.64 | -4.1656 | -309.0875 | 0
3 | (6x) 6-32 x 1/4 PHMS + #6 LW, batt. cage to rails | 55209 / 55217 | 1 | 5.3 | 5.3 | 21.78 | 55.3212 | 293.2024 | 0 | 0 | 21.78 | 55.3212 | 0 | -0.25 | -0.635 | -3.3655 | 0
2 | Assy., Battery Hull Sections | 49878
| Subtotal Weight | 8056.2 | gm | grams: diff
| Scale Weight | 8057 | gm | grams: 0.8
3 | Hull, Fwd Battery | 49817 | 1 | 3651.9 | 3651.9 | 43.539 | 110.5891 | 403860.1882 | 12257.8 | 12257.8 | 43.539 | 110.5891 | 1355578.5797 | 0 | 0 | 0 | 0
3 | Hull, Aft Battery | 49817 | 1 | 3641.2 | 3641.2 | 33.671 | 85.5243 | 311411.2268 | 12257.8 | 12257.8 | 33.671 | 85.5243 | 1048340.2549 | 0 | 0 | 0 | 0
3 | Joint Ring , battery hulls | 49818 | 1 | 685.7 | 685.7 | 38.935 | 98.8949 | 67812.2329 | 436.8 | 436.8 | 38.935 | 98.8949 | 43197.2923 | 0 | 0 | 0 | 0
3 | (8x) 6-32 x 3/4 SHCS + #6 LW, joint ring to fwd/aft batt. Hull | 55214 / 55217 | 1 | 13.5 | 13.5 | 38.625 | 98.1075 | 1324.4513 | 0 | 0 | 38.625 | 98.1075 | 0 | 0 | 0 | 0 | 0
3 | (2x) O-Ring, E70-271 | 52412 | 1 | 16.7 | 16.7 | 38.625 | 98.1075 | 1638.3953 | 7.32 | 7.32 | 38.625 | 98.1075 | 718.1469 | 0 | 0 | 0 | 0
3 | (12x) Clips, Rib | 49883 | 1 | 21.3 | 21.3 | 38.625 | 98.1075 | 2089.6898 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 213.7029 | 0
3 | (12x) Clips, Wiring | 52405 | 1 | 18.8 | 18.8 | 38.625 | 98.1075 | 1844.421 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 188.6204 | 0
3 | Oil Absorption pad, 1.9" x 10.125" + (2x) rib clips | 55527 / 49883 | 1 | 7.1 | 7.1 | 38.625 | 98.1075 | 696.5632 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 71.2343 | 0
2 | Assy., Forward Hull Sections | 49877
| Subtotal Weight | 5193.4 | gm | grams: diff
| Scale Weight | 5194 | gm | grams: 0.6
3 | Endcap, Forward | 49830 | 1 | 1918.3 | 1918.3 | 15.566 | 39.5376 | 75845.0548 | 5178.3 | 5178.3 | 15.781 | 40.0837 | 207565.6308 | 0 | 0 | 0 | 0
3 | Acoustic transducer (ITC-3013), w/ cable W36 | 52267 / 52223 | 1 | 1441.2 | 1441.2 | 9.015 | 22.8981 | 33000.7417 | 656.2 | 656.2 | 9.015 | 22.8981 | 15025.7332 | 0 | 0 | 0 | 0
3 | O-Ring, E70-238, transducer to Endcap | 52413 | 1 | 3.2 | 3.2 | 10.5 | 26.67 | 85.344 | 0 | 0 | 10.5 | 26.67 | 0 | 0 | 0 | 0 | 0
3 | (6x) Alum 1/4-20 x 3/4 SHCS + 1/4 LW, xducer to end cap | 55236 / 55252 | 1 | 15.1 | 15.1 | 10.7 | 27.178 | 410.3878 | 0 | 0 | 10.7 | 27.178 | 0 | 0 | 0 | 0 | 0
3 | Hull, Electronics | 49829 | 1 | 1701 | 1701 | 22.843 | 58.0212 | 98694.0952 | 5240.9 | 5240.9 | 22.919 | 58.2143 | 305095.1152 | 0 | 0 | 0 | 0
3 | O-Ring, E70-263, elec hull to fwd endcap | 52410 | 1 | 6.3 | 6.3 | 19.77 | 50.2158 | 316.3595 | 5.76 | 5.76 | 19.77 | 50.2158 | 289.243 | 0 | 0 | 0 | 0
3 | (8x) 6-32 x 1/2 SHCS + #6 LW, elec hull to fwd endcap | 55183 / 55217 | 1 | 10.7 | 10.7 | 19.87 | 50.4698 | 540.0269 | 0 | 0 | 19.87 | 50.4698 | 0 | 0 | 0 | 0 | 0
(pressure test above components w/ bulkhead and pressure sensor assy)
3 | Compass bracket | 49848 | 1 | 25.7 | 25.7 | 12.01 | 30.5054 | 783.9888 | 0 | 0 | 12.01 | 30.5054 | 0 | -0.35 | -0.889 | -22.8473 | 0
3 | Compass w/compass carrier PCB | 55305 / 52466 | 1 | 46.5 | 46.5 | 12.01 | 30.5054 | 1418.5011 | 0 | 0 | 12.01 | 30.5054 | 0 | 0 | 0 | 0 | 0
3 | (4x) 4-40 x 3/8 SHCS + #4 x 1/8L stand off | 55256 / 55282 | 1 | 2.8 | 2.8 | 12.01 | 30.5054 | 85.4151 | 0 | 0 | 12.01 | 30.5054 | 0 | -0.31 | -0.7874 | -2.2047 | 0
3 | (2x) 4-40 X 3/8 SHCS + #4 LW, comp mt to fwd encap | 55228 / 55233 | 1 | 1.2 | 1.2 | 10.8 | 27.432 | 32.9184 | 0 | 0 | 10.8 | 27.432 | 0 | -0.81 | -2.0574 | -2.4689 | 0
3 | Cable Assembly W9 (compass cable) | 52244 | 1 | 21.4 | 21.4 | 15.33 | 38.9382 | 833.2775 | 0 | 0 | 15.33 | 38.9382 | 0 | 0.38 | 0.9652 | 20.6553 | 0
Note on M2: Fritz Stahr:07 may 2006 Made changes to make easier and more clear to use. Added Ballast sheet with connections to both Trim sheet and Cal sheet. Added Pressure Sensor cal sheet directly here with connections to Cal sheet. Added Maintenance sheet to keep track of during build changes. Modified the Trim page to accomodate noted changes in weight of sub-assemblies. Changed top of page to interact with Ballast sheet better and to make explicit note of changes in trim weight location to get good pitch and roll gain. Russ' notes below are a good basic guide to the original intent of the sheet, but should not be taken too literally. Russ Light: (~2001?) There are two primary goals in achieve with the spreadsheet: 1) Trimming which involves getting neutral bouyancy 2) Balancing which involves achieving zero LCB-LCG separation Trimming the glider amounts to changing the weight value for the forward fairing trim weight until the Net Buoyancy value goes to zero. The Internal Oil Stroke must also be set for the desired percentage to achieve mission conditions. Thre stroke value determines the nuetral point of the VBD system and thus determines how much displacement is available for thrust and surface manuever. Of these two values the bouyancy required for the surface maneuver is the most important. The density of the surface water layer determines how much displacement in the external bladder will be required to achieve the proper antenna mast position for GPS and data telemetry. In cases of extream fresh water conditions as much as 650cc were required in the external bladder to achieve the proper surface position. Select the stroke percentage to give sufficient displacement for the surface maneuver. Adjust the forward fairing trim lead to zero the Net Buoyancy. To achieve the proper balance the LCB-LCG difference value is made to be zero. After achieving neutral buoyancy in the above procedure the center of gravity can be adjusted by changing the Pitch Mass Stroke percentage. The ideal value for this percentage is probably between 50-75%. The total throw of the pitch mass shifter is approx 15cm. Typical glider operations only require 1-3cm of movement to acheive normal glide slopes. By biasing the neutral point towards the aft more shift is provided for the surface manuever. FROM THE VMG USER'S GUIDE 10.0 Trimming Procedure This section outlines how to manipulate, or more generally, in what order to manipulate the VMG trim sheet. There is a specific trim sheet for every glider, and its maintenance and upkeep are a necessary and time-consuming task, especially with changes to the VM hardware. The goal is to set the VM Variable Buoyancy Device (VBD) for a given range of expected seawater density so that the limited range of buoyancy control is sufficient to perform its mission. After the vehicle is trimmed for neutral buoyancy at the expected density, it is then necessary to set the vehicle longitudinal pitch trim so that it floats level, with zero pitch, with sufficient remaining control over its pitch to dive and climb. The range of seawater density determines how precisely the VM must be trimmed. Density ranges greater than 0.7 practical salinity units (psu) may reduce the amount of thrust available during a dive to less than 200 grams. Such a density range will require exact laboratory trimming to avoid using some of the VBD volume for achieving neutral trim. The volume at the surface defines the minimum positive buoyancy and provides a starting point for the calculations. An estimate of the surface density (at an average depth of 0.5 to 1.0 meter) must be available for this calculation. With all the trim sheet weights and positions entered with utmost accuracy, we’re ready to set the final trim for a given density range expected. Adjustments are generally left to modifying the amount and location of external hull trim lead, which is that lead carried, or taped, to the inside of the bottom of the fairing. First we’ll set the overall buoyancy trim, then deal with the longitudinal pitch trim. Cell names in Italics represent actual cells in the spreadsheet. Set Overall Buoyancy Trim 1) Set the VBD Internal Oil Stroke to 5%. This represents the near maximum displaced volume, leaving 5% margin for trim errors and/or lighter than expected surface density. 2) Set Density to the expected surface density (e.g. 1.018) 3) Adjust the external trim lead to yield about 275 grams (+/- 10) grams of Net Buoyancy. This is required to elevate the present cellular phone/GPS antenna design to the maximum degree above the water. Future designs may take more or less: it is a function solely of the displaced volume of the antenna assembly. 4) Now set the Density to the expected seawater density at the apogee depth (e.g. 1.024). This is the density where the VM trim neutral point is defined. 5) Using the goal seek function of Microsoft EXCEL, or the hunt and peck method, determine the VBD internal oil gage setting to achieve neutral buoyancy at apogee. Force the Net Buoyancy value to zero by adjusting the value of the Internal Oil Stroke. Record this value of neutral trim in the calibration sheet (specifically the A/D value of VBD Position) 6) Observe the maximum obtainable negative buoyancy by setting the Internal Oil Stroke to 95%. This is the maximum thrust available during a dive. This should exceed that required for the mission as determined by expected currents or round trip dive times. If this maximum thrust is much more than is likely needed or desired, it might be wise to repeat the process, but leaving more positive buoyancy available for surfacing in lighter than expected surface water density (e.g. 10% or 15% versus 5%). Set Vehicle Longitudinal Pitch Trim 7) With the VBD set at the neutral trim point as determined in step #5, and the Pitch Mass Stroke set to 50%, play with the XCG position of the Trim Lead, Main Hull. Verify that there is position in the hull to accommodate the lead; this can be done in the big AutoCAD assembly drawing, #49800. It will probably be needed to place the trim lead in the most advantageous position in the fairing, and achieve the final pitch trim using the Pitch Mass Stroke, to achieve a zero Vehicle Pitch Angle. As done in step 5, record the Pitch Mass Position A/D value on the calibration sheet. Note that these two numbers, which define the neutral buoyancy and neutral pitch trim, are initial estimates only. Their values will surely be modified in the field based upon dive performance data.
Note on C18: This should be either measured or predicted denstiy the density of water the glider will encounter, e.g., OSB tank, PS surface, PS deep, WA coast deep.
Note on C19: Fritz (07May06): Minimum value for surface manuever is 150 cc so test by setting to density to lightest possible (e.g., 1.018 for PS) and then put oil reservoir to 5% to be sure it can do it.
Note on B20: Russ Light: The percentage of oil in the internal bladder. 100% stroke means the internal bladder is full. Since the spreadsheet is designed to be adjusted until neutral buoyancy is achieved this value is critical. The user must weigh the issues of how much positive buoyancy is required at the surface for surface maneuver based on surface water density. For fresh surface water conditions this can result in a lack of thrust range. Very fresh water conditions found during Puget Sound testing in Possession Sound during the spring/summer required surface buoyance values from 600-650 cc. In general the stoke is biased above 50% so that there is sufficient reserve for achieving enough positive buoyancy on the surface.
Note on E20: External Bladder Full = low AD = vehicle floats higher. External Bladder Empty = high AD = vehicle sinks lower.
Note on B27: Nominal is 70% so can pitch far forward during surface manuever, Max Stroke is 15.32 cm
Note on E27: Positive cm is aft = high AD. Negative cm is fwd = low AD.
Note on B35: Jim Osse: Itterates for roll by equating moment of battery against moment of hull VCG and hull VCB, must have itteration option ON
Note on J38: Fritz Stahr: Anything inside the pressure hull is considered to displace no volume.
Note on L38: Russ Light: In the case where the part is fabricated from a homogenous material the CG and CB are the same. In this case you will find that the CB is taken directly from the CG cell.
Note on F52: Fritz Stahr: determined by original all-up dry weight difference, including lead and tape 6/6/06, then adjusted to match all up wt. for HI
Note on G52: Fritz Stahr: set to LCG zero point
Note on J52: Fritz Stahr: set to make neutral observed in Port Susan show 0 net buoyancy
Note on L52: Fritz Stahr: adjusted to make pitch center found in PS be zero
Note on G62: Fritz Stahr: 16may06 In orig trim sheets this changed from glider to glider and appeared to be calculated. Further, this # was also pasted into the xcg for the aft fairing. Replaced with Solidworks locations for xcg/b.
Note on J62: Russ Light: Weighed fwd fairing (painted) with joint ring in air and in water: Air Wt = 6603g H20 Wt = 2595g Diff = Vol = 4008cc Using computed volumes of joint ring fwd fairing volume is = 3686cc
Note on G69: Fritz Stahr: 41.70" represents 5" long strips of weight butted up against fwd edge of fairing joint ring
Note on G81: Tom: MLW solidworks centroid-no parts included 8/8/2001: 54.15" Jim: Orig ACAD estimate 35.20" total fairing; enter only for forward fairing Fritz: trim sheets ahd been taking this from fwd fairing, I assume to represent the whole fairing, but now individual LCGs and LCBs
Note on J84: mike johnson: Volume taken from Solidworks model 49838-A
Note on J101: Volume obtained by Light and derived by twl 6/200o
Note on J107: Volume obtained by Light and derived by twl 6/200o
Note on J114: Russ Light: Volume obtained using wt in air and wt in water method using Dasaro's tank. Performed by Light on 08/04/00
Note on L114: Russ Light: From Welch e-mail on 7/3/00, derived from Solidworks dwg
Note on J115: Russ Light: Volume obtained using wt in air and wt in water method using Dasaro's tank. Performed by Light on 08/04/00
Note on L115: Russ Light: From Welch e-mail on 7/3/00, derived from Solidworks dwg
Note on J128: mike johnson: Calculated from OEM specs
Note on C147: mike johnson: Variable oil removed from weight roll up.
Note on J149: mike johnson: Volume taken from Solidworks model 49815-F. Internal features / thru holes removed to create a solid part.
Note on F175: Includes Oil
Note on J175: Jim Osse: incl vol of bladder at spec grav of 1.0
Note on B294: mike johnson: Higlighted are moving mass
Note on G95: Mike Johnson: 15 may 2008 70.18" LCG computed from solidworks model for 12" TTI antenna mast assy. Includes all components
Note on J95: mike johnson: 1/29/2008 Volume calculated with all components using solidworks models
Note on L95: mike johnson: 1/29/2008 Assume same as LCG
Note on J87: mike johnson: changed to polycarb rudder 12/9/08

## Lead Worksheet
SG lead worksheet | SG# | 139 |  | General notes re: usage
| date | 2008-12-12 |  | Lead, 1/8"x3/4"x5"≈ |  |  | ~95g
| time | noon |  | Lead, 1/4"x3/4"x5"≈ |  |  | ~171g
| who | FS/GB |  | 1pc foam double-stick≈ |  |  | ~0.5g | change/add for ocean
[8 empty] | total for ocean |  | 2703.7 | current | 2706.3 | diff | 2.6
| Wt. goal | Actual | [8 empty] | Actual
Forward Fairing | grams | grams/pcs | [8 empty] | sum grams
Nose plates (qty) | 1 | pc-1 | pc-2 | pc-3 | pc-4 | pc-5 | pc-6 | pc-7 | pc-8 | pc-9
weights | grams 1: 741.4; grams/pcs pc-1: 246; grams/pcs pc-2: 245.5; grams/pcs pc-3: 249.9 | 741.4
screws | grams 1: 23.5
At Fairing Joint Ring | grams/pcs pc-1: port/upper; grams/pcs pc-5: middle; grams/pcs pc-9: strbrd/lower
Bottom | 1217.5 |  | 170.5 | 171.5 | 171.2 | 180.2 | 182.2 | 171.2 | 170.7 |  | 1217.5
# of pcs tape | 7 |  | 1 | 1 | 1 | 1 | 1 | 1 | 1
Port side | grams 1: 233.4; grams/pcs pc-4: 42.9; grams/pcs pc-5: 95; grams/pcs pc-6: 95.5 | 233.4
# of pcs tape | grams 1: 2.5; grams/pcs pc-4: 0.5; grams/pcs pc-5: 1; grams/pcs pc-6: 1
Starboard side | grams 1: 381.1; grams/pcs pc-4: 95.1; grams/pcs pc-5: 94.7; grams/pcs pc-6: 95; grams/pcs pc-7: 96.3 | 381.1
# of pcs tape | grams 1: 4; grams/pcs pc-4: 1; grams/pcs pc-5: 1; grams/pcs pc-6: 1; grams/pcs pc-7: 1
Top | 476.1 |  |  | 96.4 | 95.1 | 95.4 | 94.3 | 94.9 |  |  | 476.1
# of pcs tape | 5 |  |  | 1 | 1 | 1 | 1 | 1
sum - tape | grams 1: 18.5 | 18.5
At 16.7" fwd of Ring | grams 1: 379.7; grams/pcs pc-4: 95.2; grams/pcs pc-5: 95.6; grams/pcs pc-6: 95; grams/pcs pc-7: 93.9 | 379.7
# of pcs tape | grams 1: 4; grams/pcs pc-4: 1; grams/pcs pc-5: 1; grams/pcs pc-6: 1; grams/pcs pc-7: 1 | 4
2706.3
[7 empty] | Scale weight of fwd fairing after lead added | 11026
Aft Fairing | grams/pcs pc-6: original
At fwd edge stern tube | grams 1: 0 | 0
# of pcs tape | grams 1: 0
Scale weight of aft fairing after lead added
Whole glider weight - BEFORE PUTTING IN TANK - DRY!!!

## Ballast
Ballasting worksheet |  | SG | 139
| Date: | 2009-07-01 18:19 | Time: | 39995.7634
| Conversions |  | Units
VBD | -4.0767 | counts/cc | cc=cm3 | gm=grams
| -0.2453 | cc/counts | counts=AtoD counts
In tank to find original volume of SG |  |  |  | New tank test 10/22/07
Tank Temp | 21.7769 | deg C | measured using YSI
Tank Salinity (meas) | 33.6 | ppt | measured using YSI
Tank Salinity (corr) | 32.4763 | ppt | 0.96656 | ratio for tank salt
Tank Density | 1.0224 | gram/cc | calculated using web-calculator
SG Mass | 51402 | grams | measured with lead est. from orig. trim sheet
SG Volume | 50276.9526 | cc | at neutral, when achieved relative to C_VBD
Neutral by interation in tank |  | -15 | cc
Relative to set C_VBD |  | 3361 | counts
(Neil's table ) | counts* | VBD rel V0, cc | Displacement
Vol max (oil outside) | 450 | 729.0567 | 51006.0093 | cc (pos. floats)
Vol neutral | 3422.1506 | (obsv in tank) | 50276.9526 | cc
Vol min (oil inside) | 3860 | -107.4027 | 50169.5499 | cc (neg. sinks)
Max movable vol. |  | 836.4594
New enviornment ballasting
SG Mass (observed) | 51385 | gm | at PS deployment
SG Vol max (obsv/est) | 50997 | cc | from PS Dec 08
SG Vol min | 50160.5406 | cc
New enivron density | 1.0275 | g/cc
Goal for thrust | 285.3533 | cc
New Vol neutral | 50445.8939 | cc | lead density | 11.296 | gm/cc
Change mass by | 493 | gm | vol lead change | 43.6438 | cc
Projected new mass | 51878 | gm
Proj. Vol. max | 51040.6438 | cc
VBD points for new environ
(Neil's table ) | counts* | VBD rel V0, cc | Displacement
Vol max | 450 | 551.1061 | 51040.6438 | cc
Vol neutral (C_VBD) | 2696.6981 | 0 | 50445.8939 | cc
Vol min | 3860 | -285.3533 | 50204.1844 | cc
Max movable vol. |  | 836.4594
Actual new mass | 51904 | gm
Est. Vol. neutral | 50514.8418 | cc
Est. C_VBD | 2415.6177 | counts
Est. Vol. max | 51040.6438 | cc

## Tank _ PS notes
Tank and Puget Sound Notes
Seaglider | 139
Date | Action
2007-08-16 | first time to Port Susan - 27 dives, high motor currents for pitch - rebuild
2007-08-28 | 2nd time to field to check on newly rebuilt mass - shifter - still high currents and climbing after only 4 dives - bring back
2007-10-22 | tank check with latest mass-shifter rebuild - much better - found new values for C_PITCH, etc. from trim out and new wts.
2008-12-10 | Back to Port Susan after new batteries:

## Maint
Maintenance sheet
SG | 139
Date | Action
2008-12-05
2008-12-09 | Finished refurb - changed both battery packs,
Changed wings for best pair from boxes with spares
Changed SBE 43F sensor with one that came back from calibration - was on SG146/8. now SN
Changed out rudder - too chipped to really use - now polycarb
2008-12-11 | Glider Refurbishment
Glider returned to Seattle Nov 08
Powered up glider and ran hardware check out - glider in working order
CF card dumped to Config folder, old data & battery files deleted, battery guage/file reset to 0,
Upgraded AAE firmware to 3.061
Updated code to Rev 66.04
Replaced TT8 (drew high current after AAE reflash)
CT removed and sent to calibration (SN 0066)
SBE43f O2 removed and sent to calibration (SN0092)
Purged excess air from VBD reservoir
10V battery replaced (standard pack)
24V battery replaced (standard pack)
Re-lubed gears with Blue Moly
Re-lubed ballscrew and rails with High Slip
Passed 100 cycle pitch test (see motor data tab)
Re-assembled pupa
Ran hardware check out - passed
Calibrated compass - 12/5/08
Epoxy filled "scratches" on forward fairing
Replaced carbon fiber rudder with polycarbonate rudder
Installed inserts for externally attachable wings
Updated spreadsheet with new batt wts, etc
Re-ballasted for Port Susan
Port Susan check out flight - 12/10-12/11
Reballast for open ocean

## Pres cal
Pressure Sensor |  | last updated | 2007-08-06 |  | Paine calibration data from sheet delivered with sensor
SG | 139 |  |  |  | Response voltages in mV |  | Excitation Voltage = 10.000V
Manufacturer | Paine | Druck |  |  | Pressure | 0 deg C |  | 30 deg C
Model # | 211-75-710-05 | PCDR 4020 | Conversion to SG counts |  | (PSIA) | Increasing | Decreasing | Increasing | Decreasing
Serial # | 214721 | xxxxxx | A/D Ref Voltage | 2.495 | 0 | -0.012 | -0.002 | -0.053 | -0.036
Calibration date | 2005-03-12 | xxxxxx | Volts/AD-count | 1.4871e-07 | 300 | 5.99 | 6.041 | 5.951 | 6.002
A/D Gain | 128 | 64 | A/D Counts/Volt | 6724335.0701 | 600 | 12.009 | 12.062 | 11.966 | 12.019
Atmos @ cal. | 14.7 | 14.x | Max counts (24 bit AD) | 16777216 | 900 | 18.014 | 18.053 | 17.977 | 18.012
Temp @ cal. | 0 | xx.x | Excitation Voltage | 5.000 | 1200 | 23.986 | 24.017 | 23.942 | 23.962
[5 empty] | 1500 | 29.949 |  | 29.893
Averages of increasing and decreasing and adjusted for Seaglider 5.000V excitation voltage
Druck calibration on APL dead-weight tester |  |  |  |  | (PSIA) | Average |  | Average
Pressure (psig) | AD counts | Variance | Notes |  | 0 | -0.0035 |  | -0.02225
From dead-weight | Mean reading | (using 40 samples) |  |  | 300 | 3.0078 |  | 2.9882
0 |  |  | After setting y-intercept with 40 samples |  | 600 | 6.0177 |  | 5.9962
100 |  |  |  |  | 900 | 9.0168 |  | 8.9973
200 |  |  |  |  | 1200 | 12.0008 |  | 11.976
400 |  |  |  |  | 1500 | 14.9745 |  | 14.9465
600
800 |  |  |  |  | 0 Deg C |  | 30 Deg C
1000 |  |  |  |  | Pressure | Output | Pressure | Output
1200 |  |  |  |  | (PSIA) | (A/D counts) | (PSIA) | (A/D counts)
1400 |  |  |  |  | 0 | -3012.5021 | 0 | -19150.9063
1600 |  |  |  |  | 300 | 2588815.2073 | 300 | 2572031.267
1500 |  |  |  |  | 600 | 5179567.0231 | 600 | 5161061.653
1400 |  |  |  |  | 900 | 7760850.9752 | 900 | 7744067.0349
1200 |  |  |  |  | 1200 | 10329224.2039 | 1200 | 10307921.5104
1000 |  |  |  |  | 1500 | 12888775.105 | 1500 | 12864675.0881
800
600 |  |  |  |  | Slope Coefficient Units (PSIA/A-D )
400 |  |  |  |  | Intercept Coefficient (PSIA)
200 |  |  |  |  | 0 Deg C |  | 30 Deg C
100 |  |  |  |  | Slope | 0.00011633 | Slope | 0.00011639
0 |  |  |  |  | Intercept | -1.1727 | Intercept | 0.60228
Slope calculations |  |  |  |  | Error Analysis 0 Deg |  |  | Error Analysis 30 Deg
Averages | Pressure (psia) | AD counts | Std deviation |  | Actual | Calculated | Error | Actual | Calculated | Error
| 0 | #DIV/0! | #DIV/0! |  | (PSIA) | (PSIA) |  | (PSIA) | (PSIA)
| 100 | #DIV/0! | #DIV/0! |  | 0 | -1.5232 |  | 0 | -1.6268
| 200 | #DIV/0! | #DIV/0! |  | 300 | 299.9789 | 0.007036% | 300 | 299.9723 | 0.009225%
| 400 | #DIV/0! | #DIV/0! |  | 600 | 601.3558 | -0.226% | 600 | 601.321 | -0.2202%
| 600 | #DIV/0! | #DIV/0! |  | 900 | 901.6313 | -0.1813% | 900 | 901.9683 | -0.2187%
| 800 | #DIV/0! | #DIV/0! |  | 1200 | 1200.4049 | -0.03374% | 1200 | 1200.3867 | -0.03222%
| 1000 | #DIV/0! | #DIV/0! |  | 1500 | 1498.1523 | 0.1232% | 1500 | 1497.9785 | 0.1348%
| 1200 | #DIV/0! | #DIV/0!
| 1400 | #DIV/0! | #DIV/0! |  | Use these values for Seaglider Calibration Coefficients
| 1500 | #DIV/0! |  |  | Slope Coef. Units (PSIG/A-D )
Intercept Coef (PSIG)
Slope | #DIV/0! |  |  |  | 0 Deg C |  |  | 30 Deg C
Intercept – psia | #DIV/0! |  |  |  | Slope | 0.00011633 |  | Slope | 0.00011639
Intercept – psig | #DIV/0! |  |  |  | Intercept | -15.8727 |  | Intercept | -14.7
[5 empty] | Depth Offset | -10.8728 |  | Depth Offset | -10.0695

## Motor Data
Motor Data
| Main Pump Motor |  |  |  | Boost Pump motor |  |  |  | Roll Motor |  |  |  | Pitch Motor |  |  |  | Amp w/full mass shifter assy. orientation - nose down, full lift
SG SN | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/main pump and oil (no load on bench) (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/micro pump (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/gear block assy. (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/gear block assy. (mA): | 100-cycle average time (sec): | 100-cycle avg. current (mA) | 100-cycle avg. current peak (mA)
139 | 60 | 31 | 25 |  | 60 | 15 | 11 | 20 | 15 | 25 | 17 | 18 | 35 | 22 | 18 | 19 | 20.2 | 163 | 222

## Material
Material Density | SG | 139
Date: | 2009-07-01 18:19 | Time: | 39995.7634
6061 T6 Alum | 2.713 | gm/cc
Lead | 11.296 | gm/cc
Naval Brass | 8.415 | gm/cc
Stainless Steel | 8.027 | gm/cc
Polypropylene | 0.9 | gm/cc
Fairing fiberglass density** | 1.6498 | gm/cc | from Port Susan 12/11/08
Zinc | 7.1 | gm/cc
PVC Density | 1.41 | gm/cc
Neoprene Density | 1.23 | gm/cc
Hydraulic Oil* Density | 0.8683 | gm/cc
PRC Wing Density | 0.672 | gm/cc
Polycarbonate | 1.21 | gm/cc
** vary to balance overall density of glider
*Mobil DTE 11M specfic gravity is 0.8683gm/cc weighed by Lehman 6/26/2001
Note on B8: Russ Light (~2000?): Using an in air and in water weight for aft fairing calculated the volume. Then calculated the density of 1.5384 from this volume and in air weight. (2006) Jim O finds this varies between ~ 1.45 & ~1.65 from DG work. Osse note 1/15/07; suggest using 1.43 as startign density for fiberglass based on average of SG 101 to SG 110 post field deployment data
Note on B13: Jim Osse: 0.672 gm/cc per Lehman test on 5/28/02 on wing#23

## Matlab
% sg_calib_constants.m
% values from spreadsheet for easy inclusion of glider params into diveplot.m, etc.
% basic glider and mission params
id_str | = | 139 | ;
mass | = | 52315.8 | ; | % from sum on trim sheet
volmax | = | 51040.6438 | ; | % from ballast worksheet final table
mission_title | = | SFC_Port_Susan | ;
rho0 | = | 1027.5 | ; | % from ballast worksheet for "new" evironment
% software limits from cal sheet:
pitch_min_cnts | = | 435 | ;
pitch_max_cnts | = | 3724 | ;
roll_min_cnts | = | 227 | ;
roll_max_cnts | = | 3931 | ;
vbd_min_cnts | = | 450 | ;
vbd_max_cnts | = | 3860 | ;
vbd_cnts_per_cc | = | -4.0767 | ;
% CT sensors cal constants
calibcomm | = | SN_0041_cal_25Apr06 | ; | % SN and cal date
t_g | = | 0.0043402 | ;
t_h | = | 0.00063043 | ;
t_i | = | 2.4021e-05 | ;
t_j | = | 2.5358e-06 | ;
c_g | = | -9.8723 | ;
c_h | = | 1.0933 | ;
c_i | = | -0.001236 | ;
c_j | = | 0.00017717 | ;
cpcor | = | -9.57e-08 | ;
ctcor | = | 3.25e-06 | ;
sbe_cond_freq_min | = | 2.9832 | ; | % kHz, from cal for 0 salinity
sbe_cond_freq_max | = | 8 | ; | % kHz, est for greater than 34.9 sal max T
sbe_temp_freq_min | = | 2.9242 | ; | % kHz, from cal for 1 deg T
sbe_temp_freq_max | = | 5.5553 | ; | % kHz, from cal for 32.5 deg T
% oxygen cal constants
comm_oxy_type | = | SBE_43f | ; | % spec SBE_43f or Aa_optode
calibcomm_oxygen | = | SN_430102_cal_25Apr06 | ; | % SN and cal date
Soc | = | 0.00023 | ;
Foffset | = | -836 | ;
A | = | -0.00025244 | ;
B | = | 3.5548e-05 | ;
C | = | -8.9129e-07 | ;
D | = | 0.036 | ;
% initial hydrodynamic model params
hd_a | = | 0.003836 | ; | % #s from SG115 in HI, July 2006
hd_b | = | 0.010078 | ;
hd_c | = | 9.8541e-06 | ;