## Cal
Calibration Sheet for SEAGLIDER |  | Serial No. | 114
Date: | 2008-06-24 10:57 |  | (master to all sheets)
Time: | 39623.4559
Software Revision | Rev 65.03/1099
Pitch Mass |  | Last updated | 2008-05-27
Comments
| Headroom, fwd | 25 | AD counts
| Headroom, aft | 25 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | cm travel*
Minimum (full forward) | 385 | 410 | -9.5241
Maximum (full aft) | 3724 | 3699 | 5.6053
Stroke Length |  |  | 15.1294
Pitch Center | (as defined by $C_PITCH) | 2480.4575 | 0
(15.32 cm full stroke)
Conversion constant | 0.0046 | cm per AD | ($PITCH_CNV)
(conversion inverse) | 217.39 | AD per cm
*equation | cm = (AD –center)* conversion factor (cm/AD)
(-) cm is fwd of center, (+) cm is aft of center
Roll Mass |  | Last updated | 2008-06-23
Comments
| Headroom | 150 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | degrees roll*
Full roll to port | 10 | 160 | -50.6033
Full roll to starboard | 4090 | 3940 | 56.2573
Dive Roll Center |  | 1950 | ($C_ROLL_DIVE)
Climb Roll Center |  | 1980 | ($C_ROLL_CLIMB)
Conversion constant | 0.02827 | degree per AD | ($ROLL_CONV)
(conversion inverse) | 35.37 | AD per degree
*Equation: | deg=(AD-center) * conversion factor
Variable Buoyancy Drive (VBD) |  | Last updated | 2008-05-27
Comments
LeDuc Pump SN | NA | (short shaft SN)
Boost Pump SN | NA
| Headroom | 100 | AD counts
| Hardware Limit (counts) | Software Limit (counts) | cm3 (cc) oil*
$VBD_Min (Ext Bladder full) | 320 | 420 | 567.4601
$VBD_Max (Ext Bladder empty) | 4049 | 3949 | -298.1895
Total movable oil volume |  |  | 865.6496
Neutral Trim (from Ballast sheet) | ($C_VBD for neutral) | 2733.3688 | (nomially 0 at spec rho)
Conversion constant | -0.2453 | cc per AD | $VBD_CNV
(conversion inverse) | -4.0767 | AD per cc
*equation: | cc oil=(AD@neutral trim)* conv. factor (cc/counts)
(1 cc oil = 1.025 g buoyancy)
Pressure Sensor |  | Last updated | 2008-03-11
(all values fed from press-cal worksheet including date) |  | Comments
Values from Pres Cal sht
Manufacturer | Paine
Model # | 211-75-710-05
Serial # | 214718
Calibration date | 37782
A/D Gain | 128 |  | PSI = Slope *AD_Count + Y_Intercept
Atmos @ cal. | 14.7
Temp @ cal. | 0 | Psig/AD count
Calibrated Slope | 0.00011638 | (either PresCal-B49 or PreCal-I56)
Cal. Y-Intercept | -6.1678 | (Changed each time sensor set for “sealevel”)
Depth Offset | -4.2249 | Meters
| Conversion Factor | 0.685 psig/meter
Internal Pressure |  | Last updated | 2008-06-18
Initial psi: | 14.75 | End psi: | 9.74
Compass |  | Last updated | 2008-06-02
Comments
Manufactuer | Sparton
Model number | SP3003D
Serial Number | A691
Compass SW version # | 7.4
Compass carier PCB | PN 52281 rev 1.2
Sand Point Calibration Date | 2008-05-23
TCM2MAT file creation date | 2008-06-02
Main Board |  | Last updated | 2008-05-27
Comments
Model # | 52290 | (52290=normal glider)
Mainboard Revision | B.4
Mainboard Serial Number | 52290-009
Computer Type | TT8 | Onset Computers
Computer Serial Number | 903071
Disk system | CF8V2 | Persistor Corp.
Serial Number | 1927
Compact Flash Card Manu. | Silicon Systems
CF card size | 256 MB
CF card serial/ID # | SSD-C25MI-3005
18343
Watchdog (PIC) setting (min) | 10 min | (10min = dips 2&4 on)
OEM Board & Sub-systems |  | Last updated | 2008-05-27
Comments
GPS Manufacturer | Garmin
GPS Model Number | 15H-W
GPS Serial Number | 27701226
GPSI Antenna SN | 0081 | (fab by Ocean Eng Services)
DC-DC converter SN | 11009A545
Transducer Manu. | ITC
Transducer Model | 3013
Transducer SN | 3734
Acoustic Transponder |  | Last updated | 2008-05-27
Comments
Manufacturer | AAE
Model | 955 | (with altimeter function) | (5414-5000?)
Main-board SN | 0036
Sub-board SN | 286
AAE software ver # | Ver 3.5
Interrogate Frequency | 12 | KHz
Reply Frequency | 11.5 | KHz
DIP switch settings | 9/6 | (sw-1/sw-2)
Iridium Modem |  | Last updated | 2008-06-23
Comments
Model | 9522A
IMEI Number | 300224010030220
modem SN: | 4F3XW
SIM Card Owner | UH | phone #
SIM Card SN | 89881 69414 00043 9602
SBE C and T Sensors |  | Last updated | 2008-06-04
Comments
Sail SN | 38
Temp PCB SN | 38
Cond PCB SN | 38
SBE Calibration date: | 2007-12-18
Conductivity Calibration Coefficients |  | Temperature Calibration Coefficients
g = | -10.0024 | g = | 0.0044131
h = | 1.1227 | h = | 0.00065298
i = | -0.001173 | i = | 2.7242e-05
j = | 0.00011679 | j = | 2.9467e-06
Cpcor (nom) = | -9.57e-08 | f0 = | 1000
Ctcor (nom) = | 3.25e-06
Optical Sensor(s) |  | Last updated | 2008-06-04
Manufacturer | WET Labs
Model | (BB2F-VMG)
Serial Number | 451
Dissolved Oxygen Sensor |  | Last updated | 2008-06-04
Manufacturer | SeaBird | Sea-Bird 43f coefficients | cal date: 23-Jan-08
Model | 43f | Soc = | 0.00025391
Serial Number | 98 | Boc =
Tcor =
Manufacturer | Aanderaa | Pcor =
Model | Optode | Foffset= | -806.5054
Serial Number | xxx | A= | -0.0017638
|  | B= | 0.00017276
|  | C= | -3.4856e-06
|  | E= | 0.036

## Trim
| SEAGLIDER - Trim and Balance |  | Date: | 2008-06-24 10:57
| SG | 114 | Time: | 39623.4559 | [7 empty] | Spreadsheet Usage Notes from Russ Light
Comments:
[11 empty] | Fwd Fairing assy scale wt: | 8274
|  | Enter data only where font is RED | [8 empty] | Aft fairing assy scale wt: | 4780.2
| Density Trim | [9 empty] | Fairing screws scale wt: | 39
| Total Weight (in air), summed items | 52448.2 | gm | [7 empty] | Pupa assy w/lead scale wt: | 39353
| Total Scale Weight | 52446.2 | gm | wt from sum of assy scale wts | [6 empty] | Total glider wt by assy scale wts: | 52446.2
| Weight difference/"error" | -2 | gm
| Variable wt, including nose-wt-plates | 3841 | gm
| Displaced Volume, measured in tank | 51238.7321 | cc | diff | see Ballast worksheet for more volume details
| Displaced Volume, summed items | 51044.4769 | cc | -194.2552
| Water Density | 1.0275 | gm/cc
| Net Buoyancy (relative to neutral, positive floats, no thrust) | 0 | cc |  | predicted using this spreadsheet, actual may vary!
| Internal Oil Stroke (100% is all inside, ext. bladder empty) | 65.55% |  | 2733.3688 | A/D counts | (input %, not counts)
[6 empty] | ($PITCH_VBD_SHIFT) | 0.00167 | cm per cc
Pitch & Roll Trim
| Longitudinal CG | 89.5463 | cm | 35.2544 | inches
| Longitudinal CB | 89.5463 | cm | 35.2544 | inches
| LCB-LCG Separation | 6.035e-08 | cm
| Pitch-mass Stroke (100% is fully aft, nominal is 70%) | 62.95% |  | 2480.4575 | A/D counts | (input %, not counts)
| Pitch-mass (batt) Location | 9.6441 | cm
| Vehicle Pitch Angle (positive is nose up) | 5.6531e-06 | deg | (goal is change of 15-20 deg per cm travel, use goal-seek on Location by +/-1cm to test)
| Vertical CG (VCG) | -0.55156 | cm | (goal is ~ -0.56 from prior builds)
| Computed VCB | 0.060113 | cm | (goal is ~0.05 from prior builds)
| VCB-VCG Separation | 0.61167 | cm | (goal is ~0.61 from prior builds)
| Roll-mass (batt) Angle | 0 | deg
| Vehicle Roll Angle | #N/A | deg | (goal is 20 deg Vehicle roll per 40 deg Battery Roll, but eqn had circular reference, now repl w/"20")
[5 empty] | Total |  |  | Moment |  |  |  |  | Moment |  |  | Moment | Moment
Assy. |  | Part # | Qty | Weight | Weight | LCG | LCG | LCG | Volume | Total Vol | LCB | LCB | LCB | VCG&B | VCG&B | VCG | VCB
Assy. Level | Description | (or Assy Wt) |  | grams | grams | inch | cm | gm-cm | cc | cc | inch | cm | cc-cm | inch | cm | gm-cm | cc-cm
Complete fairing & antenna
| Subtotal (summed) Weight | 13053.9 | gm | grams: diff
| Scale Weight | 13054 | gm | grams: 0.1
| Summed volume | 8187.3464 | cc
|  |  |  | sum-diff's | 45.8
Complete pupa
| Subtotal (summed) Weight | 36251.3 | gm | grams: diff
| Scale Weight | 36297 | gm | grams: 45.7
| Summed volume | 42857.1305 | cc
|  | Total weights and volume - Flying SG - without adjustment |  | grams: 52402.5; gm-cm: 4680455.377; cc: 51044.4769; cc-cm: 4570842.8416; gm-cm: -28928.2904; cc-cm: 3068.415
|  | Adjustment to reflect trims found in tank & Port Susan |  |  | 45.7 | 138.5762 | 351.9835 | 16085.6459 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Total weights and volume - Flying SG - adjusted, reported above |  | grams: 52448.2; gm-cm: 4696541.0229; cc: 51044.4769; cc-cm: 4570842.8416; gm-cm: -28928.2904; cc-cm: 3068.415
0 | Flying Seaglider | 49801
1 | (8x) 10-32 x 3/8 FHMS, joint ring to aft fairing | 55240 | 1 | 11.7 | 11.7 | 46 | 116.84 | 1367.028 | 1.4576 | 1.4576 | 46 | 116.84 | 170.3037 | 0 | 0 | 0 | 0
1 | (8x) 1/4-28 x 1/2 FHMS, Aft fairing to endcap | 55294 | 1 | 27.3 | 27.3 | 49.5 | 125.73 | 3432.429 | 3.401 | 3.401 | 49.5 | 125.73 | 427.6104 | 0 | 0 | 0 | 0
1 | Assy, Forward Fairing | 55494
| Subtotal Weight | 8273.7 | gm | grams: diff
| Scale Weight | 8274 | gm | grams: 0.3
2 | Forward Fairing | 49897 | 1 | 5550.8 | 5550.8 | 28.22 | 71.6788 | 397874.683 | 3398.8546 | 3398.8546 | 28.22 | 71.6788 | 243625.8207 | 0 | 0 | 0 | 0
2 | Joint Ring, Fwd to Aft Fairing (bonded to fwd fairing before paint; Long ring = 875g, short ring 688g ) | 49834 | 1 | 688 | 688 | 45.85 | 116.459 | 80123.792 | 253.5938 | 253.5938 | 45.85 | 116.459 | 29533.2812 | 0 | 0 | 0 | 0
2 | Nose weight (bonded to fwd fairing after paint) | 49844 | 1 | 1252.2 | 1252.2 | 3.715 | 9.4361 | 11815.8844 | 148.8057 | 148.8057 | 3.715 | 9.4361 | 1404.1455 | 0 | 0 | 0 | 0
2 | Nose weight plates | 55448 | 1 | 759 | 759 | 3.715 | 9.4361 | 7161.9999 | 90.1961 | 90.1961 | 3.715 | 9.4361 | 851.0992 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 2" SHCS + 1/4 LW brass, nose plate to nose base (4 - 5 plates) | 55186 / 55253 | 0 | 28.5 | 0 | 3.715 | 9.4361 | 0 | 3.5505 | 0 | 3.715 | 9.4361 | 0 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 1.5" SHCS + 1/4 LW brass, nose plate to nose base (2 -3 plates) | 55185 / 55253 | 1 | 23.7 | 23.7 | 3.715 | 9.4361 | 223.6356 | 2.9525 | 2.9525 | 3.715 | 9.4361 | 27.8604 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x .75" SHCS + 1/4 LW brass, nose plate to nose base (1 plate) | 55473 / 55253 | 0 | 15.5 | 0 | 3.715 | 9.4361 | 0 | 1.931 | 0 | 3.715 | 9.4361 | 0 | 0 | 0 | 0 | 0
1 | Assy., Lead Ballast
| Subtotal Weight | 3058.3 | gm | grams: diff
| Scale Weight | 0 | gm | grams: -3058.3
2 | Trim Lead .600" aft of pupa joint ring - bottom (5" strip) |  | 1 | 1452.8 | 1452.8 | 42.077 | 106.8756 | 155268.8426 | 128.6119 | 128.6119 | 42.077 | 106.8756 | 13745.4712 | -5.16 | -13.1064 | -19040.9779 | -1685.639
2 | Trim Lead .600" aft of pupa joint ring - port side (5" strip) |  | 1 | 392.2 | 392.2 | 42.077 | 106.8756 | 41916.6025 | 34.7203 | 34.7203 | 42.077 | 106.8756 | 3710.7474 | 0 | 0 | 0 | 0
2 | Trim Lead .600" aft of pupa joint ring - starboard side (5" strip) |  | 1 | 487.1 | 487.1 | 42.077 | 106.8756 | 52059.095 | 43.1215 | 43.1215 | 42.077 | 106.8756 | 4608.6309 | 0 | 0 | 0 | 0
2 | Trim Lead .600" aft of pupa joint ring - top (5" strip) |  | 1 | 454.9 | 454.9 | 42.077 | 106.8756 | 48617.7013 | 40.2709 | 40.2709 | 42.077 | 106.8756 | 4303.975 | 5.1 | 12.954 | 5892.7746 | 521.6691
2 | Trim Lead tape (~1g per 5" lead strip) |  | 1 | 22 | 22 | 42.077 | 106.8756 | 2351.2628 | 15.6028 | 15.6028 | 42.077 | 106.8756 | 1667.5622 | 5.1 | 12.954 | 284.988 | 202.1191
2 | EPDM base (5"x31") + (2x) 30" tape | not in BOM | 1 | 218.2 | 218.2 | 42.077 | 106.8756 | 23320.2516 | 156.5729 | 156.5729 | 42.077 | 106.8756 | 16733.82 | 0 | 0 | 0 | 0
2 | (2) Straps - cut to length (start at 34g, subtract trimmed) | not in BOM | 1 | 31.1 | 31.1 | 42.077 | 106.8756 | 3323.8305 | 27.02 | 27.02 | 42.077 | 106.8756 | 2887.7763 | 0 | 0 | 0 | 0
2 | Trim Lead 1.000" fwd of bulkhead - bottom (5" strip) |  | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 4.035 | 10.2489 | 0 | 0
2 | Trim Lead tape (~1g per 5" lead strip) |  | 0 | 0 | 0 | 42.077 | 106.8756 | 0 | 0 | 0 | 42.077 | 106.8756 | 0 | 5.1 | 12.954 | 0 | 0
2 | EPDM base (5"x3") | not in BOM | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 4 | 10.16 | 0 | 0
2 | Strap - cut to length (start at 17g, subtract trimmed) | not in BOM | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 0 | 0 | 22.45 | 57.023 | 0 | 0 | 0 | 0 | 0
| lead goal | 0
| currently at | 3058.3 | -3058.3 | grams: diff
1 | Assy., Aft Fairing | 55497
| Subtotal Weight | 4780.2 | gm | grams: diff
| Scale Weight | 4780 | gm | grams: -0.2
2 | Aft Fairing | 49896 | 1 | 2211.8 | 2211.8 | 54.37 | 138.0998 | 305449.1376 | 1354.3249 | 1354.3249 | 54.37 | 138.0998 | 187031.9978 | -0.02 | -0.0508 | -112.3594 | -68.7997
2 | Wing, Starboard | 49838 | 1 | 670.2 | 670.2 | 50.942 | 129.3927 | 86718.9741 | 695 | 695 | 50.942 | 129.3927 | 89927.9126 | 0 | 0 | 0 | 0
2 | Wing, Port | 49838 | 1 | 687.2 | 687.2 | 50.942 | 129.3927 | 88918.6497 | 695 | 695 | 50.942 | 129.3927 | 89927.9126 | 0 | 0 | 0 | 0
2 | (16x) 6-32 x 3/8 FHMS, wings to fairing |  | 1 | 10.5 | 10.5 | 50 | 127 | 1333.5 | 1.3081 | 1.3081 | 50 | 127 | 166.1268 | 0 | 0 | 0 | 0
2 | Rudder | 49839 | 1 | 120.3 | 120.3 | 67.578 | 171.6481 | 20649.2688 | 99.4215 | 99.4215 | 67.578 | 171.6481 | 17065.5114 | 0 | 0 | 0 | 0
2 | (2x) 1/4-20 x 2" FHMS, rudder to rudder shoe |  | 1 | 20.3 | 20.3 | 67.75 | 172.085 | 3493.3255 | 2.529 | 2.529 | 67.75 | 172.085 | 435.1969 | 0 | 0 | 0 | 0
2 | Trim lead at rudder tube (5" strip) | 52442 / 55487 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | -1.65 | -4.191 | 0 | 0
2 | Trim Lead Tape | 55298 / 55406 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | 0 | 0 | 62.2 | 157.988 | 0 | -1.65 | -4.191 | 0 | 0
2 | Assy., Antenna | 49841
| Subtotal Weight | 483.1 | gm | grams: diff
| Scale Weight | 483.1 | gm | grams: 0
3 | Antenna Assembly GPSI (All parts) |  | 1 | 483.1 | 483.1 | 81.305 | 206.5147 | 99767.2516 | 355.5 | 355.5 | 81.305 | 206.5147 | 73415.9759 | 0 | 0 | 0 | 0
3 | Cap (DGO protector during build - do not weigh) | 52481
2 | Assy., Aanderaa Optode Oxygen Sensor | 55841
| Subtotal Weight | 0 | gm | grams: diff
| Scale Weight | 0 | gm | grams: 0
3 | O2 Sensor Assy. (Aanderaa optode, mount base, SS hose clamp, (2x) 8-32 x 5/8 FHMS, (2x) 8-32 nylok) | 52220 / 52485 / 52486 / 55239 / 55321 | 1 | 0 | 0 | 56.37 | 143.1798 | 0 | 81.1 | 81.1 | 56.37 | 143.1798 | 11611.8818 | 2.5 | 6.35 | 0 | 514.985
3 | O2 Sensor Cable Assembly | 55469 | 1 | 0 | 0 | 50 | 127 | 0 | 29 | 29 | 50 | 127 | 3683 | 2 | 5.08 | 0 | 147.32
2 | Assy., SBE 43f Oxygen Sensor | 55482
| Subtotal Weight | 342.3 | gm | grams: diff
| Scale Weight | 342.3 | gm | grams: 0
3 | O2 Sensor Assy. (O2 sensor, sensor flow tube, (2x) 6-32 x 5/8 SHCS - flow tube to sensor, (2x) O2 clamp, (4x) 6-32 x 1/2 FHMS - clamp to aft fairing) | 52356 / 52501 / 55214 / 52311 / 55210 | 1 | 283.9 | 283.9 | 54.39 | 138.1506 | 39220.9553 | 166.7 | 166.7 | 54.39 | 138.1506 | 23029.705 | 2.5 | 6.35 | 1802.765 | 1058.545
3 | O2 Sensor Cable Assembly | 52236 | 1 | 58.4 | 58.4 | 50 | 127 | 7416.8 | 29 | 29 | 50 | 127 | 3683 | 2 | 5.08 | 296.672 | 147.32
3 | Plug assy, SBE 43f plenum | 55518 | 1 | bar: grams
2 | Assy., WET Labs Optical - forward hole | 55521
| Subtotal Weight | 234.5 | gm | grams: diff
| Scale Weight | 234.5 | gm | grams: 0
3 | Wet Labs sensor Assy. (WL BB2F-VMG sensor, (2x) sensor clamp, (2x) 8-32 x 5/8 FHMS - sensor clamp to fairing) | 52296 / 52312 / 55245 | 1 | 234.5 | 234.5 | 57.7 | 146.558 | 34367.851 | 178.84 | 178.84 | 57.7 | 146.558 | 26210.4327 | -1.237 | -3.142 | -736.7943 | -561.9117
3 | Cover, WL sensor used in shipping | 55520 | 1 | bar: grams
2 | Assy., WET Labs Optical - aft hole | 55521
| Subtotal Weight | 0 | gm | grams: diff
| Scale Weight | 0 | gm | grams: 0
3 | Wet Labs sensor Assy. (WL Triplet sensor, (2x) Wet labs sensor clamp, (2x) 8-32 x 5/8 FHMS - sensor clamp to fairing) | 55381/ 52312 / 55245 | 1 | 0 | 0 | 60.55 | 153.797 | 0 | 159.3 | 159.3 | 60.55 | 153.797 | 24499.8621 | -1.237 | -3.142 | 0 | -500.5174
3 | Cover, WL sensor used in shipping | 55520 | 1 | bar: grams
1 | Assy., Pupa | 55486
| Subtotal Weight | 36251.3 | gm | grams: diff
| Scale Weight | 36297 | gm | grams: 45.7
| Complete pupae volume - sum this sheet | 42825.0014 | cc
2 | Position B (Opt O2): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 0 | 0 | 51.75 | 131.445 | 0 | 0 | 0 | 0 | 0
2 | Position C (WL-1): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 0 | 0 | 51.75 | 131.445 | 0 | -1.7 | -4.318 | 0 | 0
2 | Position D (SBE O2): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 0 | 0 | 51.75 | 131.445 | 0 | -2.3 | -5.842 | 0 | 0
2 | Position E (WL-2): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 0 | 0 | 51.75 | 131.445 | 0 | -1.7 | -4.318 | 0 | 0
2 | Position F (WL-3): IE55 dummy plug | 55309 | 1 | 0 | 0 | 51.75 | 131.445 | 0 | 3.7 | 3.7 | 51.75 | 131.445 | 486.3465 | 1.7 | 4.318 | 0 | 15.9766
2 | O-Ring, E70-270, Endcap to Batt Hull | 52411 | 1 | 8 | 8 | 48.55 | 123.317 | 986.536 | 7.13 | 7.13 | 48.55 | 123.317 | 879.2502 | 0 | 0 | 0 | 0
2 | (2x) 6-32 x 1/2 SHCS w/ LW + (5x) 6-32 x 5/8 with zincs, endcap to battery hull | 55183 / 55217 / 55214 / 49884 | 1 | 13.2 | 13.2 | 48.64 | 123.5456 | 1630.8019 | 1.6444 | 1.6444 | 48.64 | 123.5456 | 203.1646 | 0 | 0 | 0 | 0
2 | O-Ring, E70-270, Bulkhead to batt. hull | 52411 | 1 | 8 | 8 | 28.72 | 72.9488 | 583.5904 | 7.13 | 7.13 | 28.72 | 72.9488 | 520.1249 | 0 | 0 | 0 | 0
2 | (2x) 6-32 x 1/2 SHCS w/ LW + (6x) 6-32 x 5/8 with zincs, bulkhead to battery hull | 55183 / 55217 / 55214 / 49884 | 1 | 15.3 | 15.3 | 25.75 | 65.405 | 1000.6965 | 1.9061 | 1.9061 | 25.75 | 65.405 | 124.6663 | 0 | 0 | 0 | 0
2 | (4x) 6-32 x 1/2 SHCS + #6 LW, rails to bulkhead | 55183 / 55217 | 1 | 5.3 | 5.3 | 23.41 | 59.4614 | 315.1454 | 0 | 0 | 23.41 | 59.4614 | 0 | 0 | 0 | 0 | 0
2 | O-Ring, E70-263, elec hull to bulkhead | 52410 | 1 | 6.6 | 6.6 | 26.13 | 66.3702 | 438.0433 | 5.76 | 5.76 | 26.13 | 66.3702 | 382.2924 | 0 | 0 | 0 | 0
2 | (8x) 6-32 x 1/2 SHCS + #6 LW, elec hull to bulkhead | 55183 / 55217 | 1 | 10.7 | 10.7 | 25.92 | 65.8368 | 704.4538 | 0 | 0 | 25.92 | 65.8368 | 0 | 0 | 0 | 0 | 0
2 | Assy., Aft Endcap | 49891
| Subtotal Weight | 7288.9 | gm | grams: diff
| Scale Weight | 7355 | gm | grams: 66.1
3 | Hydraulic oil: End Cap Without Oil | 5637 | gm
| End Cap With Oil | 6892 | gm
| Oil Added To System | 1255 | gm
| Volume of Oil In System | 1445.353 | cc
3 | Assy., Hydraulics drive system | 55480
| Subtotal Weight | 4004.8 | gm | grams: diff
| Scale Weight | 4004.2 | gm | grams: -0.6
4 | Endcap, Aft | 49815 | 1 | 2028.1 | 2028.1 | 50.014 | 127.0356 | 257640.8192 | 2885 | 2885 | 49.752 | 126.3701 | 364577.6808 | 0 | 0 | 0 | 0
4 | (2x) Roll pin, 5/32 OD x 3/8, Pump anti rotation | 55248 | 1 | 1 | 1 | 48.41 | 122.9614 | 122.9614 | 0 | 0 | 48.41 | 122.9614 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/8 NPT x 1/4 tube, alum | 52432 | 1 | 8.4 | 8.4 | 49.3 | 125.222 | 1051.8648 | 0 | 0 | 49.3 | 125.222 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/8 NPT x 1/4 tube, SS | 52429 | 1 | 23.3 | 23.3 | 49.3 | 125.222 | 2917.6726 | 0 | 0 | 49.3 | 125.222 | 0 | 0.5 | 1.27 | 29.591 | 0
4 | Assy, Solenoid Valve, Skinner | 52254 | 1 | 610.2 | 610.2 | 49.02 | 124.5108 | 75976.4902 | 0 | 0 | 49.02 | 124.5108 | 0 | 0.6 | 1.524 | 929.9448 | 0
4 | Adaptor, Male, 1/8 NPT x 1/4 M, SS | 52430 | 1 | 10.3 | 10.3 | 50.43 | 128.0922 | 1319.3497 | 0 | 0 | 50.43 | 128.0922 | 0 | 0.5 | 1.27 | 13.081 | 0
4 | Branch Tee, Male, 1/4 tube x 1/8 NPT x 1/4 tube, alum | 52433 | 1 | 19 | 19 | 48.79 | 123.9266 | 2354.6054 | 0 | 0 | 48.79 | 123.9266 | 0 | -0.5 | -1.27 | -24.13 | 0
4 | Tubing, 1/4 DIA x 8" | 52419 | 1 | 3.3 | 3.3 | 47 | 119.38 | 393.954 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Tubing, 1/4 Dia x 6.0" | 52419 | 1 | 2.7 | 2.7 | 47 | 119.38 | 322.326 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Check Valve, Kepner 204 A-1 | 52448 | 1 | 22.7 | 22.7 | 49.1 | 124.714 | 2831.0078 | 0 | 0 | 49.1 | 124.714 | 0 | -0.5 | -1.27 | -28.829 | 0
4 | Union Elbow, Male, 1/4 tube x 1/4 tube, alum | 52435 | 1 | 16.4 | 16.4 | 50.43 | 128.0922 | 2100.7121 | 0 | 0 | 50.43 | 128.0922 | 0 | -0.5 | -1.27 | -20.828 | 0
4 | Tubing, 1/4 DIA x 9.625" | 52419 | 1 | 3.9 | 3.9 | 47 | 119.38 | 465.582 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Run Tee, Male,1/4 tube x 1/4 tube x 1/8 NPT, alum | 52428 | 1 | 19 | 19 | 49.3 | 125.222 | 2379.218 | 0 | 0 | 49.3 | 125.222 | 0 | 0 | 0 | 0 | 0
4 | Tubing, 1/4 Dia x 4.375" | 52419 | 1 | 2 | 2 | 47 | 119.38 | 238.76 | 0 | 0 | 47 | 119.38 | 0 | 0 | 0 | 0 | 0
4 | Seal Washer | 52350 | 1 | 1 | 1 | 50.9 | 129.286 | 129.286 | 0 | 0 | 50.9 | 129.286 | 0 | 0 | 0 | 0 | 0
4 | LeDuc Output Screw | 52349 | 1 | 7.6 | 7.6 | 51.04 | 129.6416 | 985.2762 | 0 | 0 | 51.04 | 129.6416 | 0 | 0 | 0 | 0 | 0
4 | Hydraulic Pump, LeDuc | 52348 | 1 | 319.9 | 319.9 | 49.58 | 125.9332 | 40286.0307 | 0 | 0 | 49.58 | 125.9332 | 0 | 0 | 0 | 0 | 0
4 | Hydraulic endcap | 49803 | 1 | 93.7 | 93.7 | 47.72 | 121.2088 | 11357.2646 | 0 | 0 | 47.72 | 121.2088 | 0 | 0 | 0 | 0 | 0
4 | Ball Bearing, 8mm ID x 16mm OD | 52459 | 1 | 3.1 | 3.1 | 47.69 | 121.1326 | 375.5111 | 0 | 0 | 47.69 | 121.1326 | 0 | 0 | 0 | 0 | 0
4 | O-Ring, N70-032, hydraulic endcap | 52407 | 1 | 0.5 | 0.5 | 48.23 | 122.5042 | 61.2521 | 0 | 0 | 48.23 | 122.5042 | 0 | 0 | 0 | 0 | 0
4 | Connector, Male, 1/16 NPT x 1/8 tube, brass | 52436 | 1 | 16 | 16 | 47.75 | 121.285 | 1940.56 | 0 | 0 | 47.75 | 121.285 | 0 | 1.15 | 2.921 | 46.736 | 0
4 | Tubing, 1/8 Dia. X 10.5" | 52418 | 1 | 1.2 | 1.2 | 45.17 | 114.7318 | 137.6782 | 0 | 0 | 45.17 | 114.7318 | 0 | 0 | 0 | 0 | 0
4 | Magnetic Seal | 52353 | 1 | 13.8 | 13.8 | 47.69 | 121.1326 | 1671.6299 | 0 | 0 | 47.69 | 121.1326 | 0 | 0 | 0 | 0 | 0
4 | (6x) 6-32 x 1/2 SHCS + #6 LW, hydraulic encap to encap | 55183 / 55217 | 1 | 8 | 8 | 48.12 | 122.2248 | 977.7984 | 0 | 0 | 48.12 | 122.2248 | 0 | 0 | 0 | 0 | 0
4 | Lee Plug Set | 52445 / 52446 | 1 | 0.4 | 0.4 | 51.512 | 130.8405 | 52.3362 | 0 | 0 | 51.512 | 130.8405 | 0 | -0.5 | -1.27 | -0.508 | 0
4 | Check Valve (altered item Kepner 1106 A-1-1) + O-Ring (N70 3-906) | 49852 / 52408 | 1 | 47.7 | 47.7 | 52.47 | 133.2738 | 6357.1603 | 0 | 0 | 52.47 | 133.2738 |  | 0 | 0 | 0 | 0
4 | Bladder, External | 52441 | 1 | 112.1 | 874.3744 | 54.613 | 138.717 | 121290.6076 | 989.9929 | 989.9929 | 54.613 | 138.717 | 137328.86 | 0 | 0 | 0 | 0
4 | Nut, Bladder | 49802 | 1 | 108.6 | 108.6 | 51.6 | 131.064 | 14233.5504 | 40.0295 | 40.0295 | 51.6 | 131.064 | 5246.4248 | 0 | 0 | 0 | 0
4 | Position A (Comms): IE55 w o-ring | 52160 | 1 | 10.6 | 10.6 | 51.512 | 130.8405 | 1386.9091 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | 1.7 | 4.318 | 45.7708 | 3.4242
4 | Position B (Opt O2): Alum port plug | 52160 | 1 | 4.1 | 4.1 | 51.512 | 130.8405 | 536.446 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | 0 | 0 | 0 | 0
4 | Position C (WL-1): IE55 w o-ring | 52160 | 1 | 11.6 | 11.6 | 51.512 | 130.8405 | 1517.7496 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | -1.7 | -4.318 | -50.0888 | -3.4242
4 | Position D (SBE O2): IE55 w o-ring | 52160 | 1 | 11.6 | 11.6 | 51.512 | 130.8405 | 1517.7496 | 0.793 | 0.793 | 51.512 | 130.8405 | 103.7565 | -1.7 | -4.318 | -50.0888 | -3.4242
4 | Position E (WL-2): Alum port plug | 52160 | 1 | 4.1 | 4.1 | 51.512 | 130.8405 | 536.446 | 1.3 | 1.3 | 51.512 | 130.8405 | 170.0926 | 0 | 0 | 0 | 0
4 | Position F (WL-3): Alum port plug | 52160 | 1 | 4.1 | 4.1 | 51.512 | 130.8405 | 536.446 | 1.3 | 1.3 | 51.512 | 130.8405 | 170.0926 | 1.7 | 4.318 | 17.7038 | 5.6134
4 | Iridium/GPS Ant., conn/o-ring/ CA W5, J2 | 52241 | 1 | 48.7 | 48.7 | 51.512 | 130.8405 | 6371.9314 | 6.8 | 6.8 | 51.512 | 130.8405 | 889.7153 | -2.3 | -5.842 | -284.5054 | -39.7256
4 | Vacuum/pressure relief valve (deep sea power) | 52449 | 1 | 11.6 | 11.6 | 51.512 | 130.8405 | 1517.7496 | 2.7 | 2.7 | 51.512 | 130.8405 | 353.2693 | 2.3 | 5.842 | 67.7672 | 15.7734
4 | Pump Plate, Inner | 49804 | 1 | 40.1 | 40.1 | 47.3 | 120.142 | 4817.6942 | 0 | 0 | 47.3 | 120.142 | 0 | 0.75 | 1.905 | 76.3905 | 0
4 | Assy, Main pump motor w/ CA W8,W18 | 52253 | 1 | 129.5 | 129.5 | 48.5 | 123.19 | 15953.105 | 0 | 0 | 48.5 | 123.19 | 0 | 1.8 | 4.572 | 592.074 | 0
4 | (3x) M2 x 6 FHMS (motor to interface plate) | 55203 | 1 | 0.4 | 0.4 | 47.3 | 120.142 | 48.0568 | 0 | 0 | 47.3 | 120.142 | 0 | 1.8 | 4.572 | 1.8288 | 0
4 | Bore Reducer | 52351 | 1 | 0.7 | 0.7 | 46.77 | 118.7958 | 83.1571 | 0 | 0 | 46.77 | 118.7958 | 0 | 0 | 0 | 0 | 0
4 | Pulley, LeDuc pump motor | 52352 | 1 | 5.7 | 5.7 | 46.77 | 118.7958 | 677.1361 | 0 | 0 | 46.77 | 118.7958 | 0 | 1.85 | 4.699 | 26.7843 | 0
4 | (6x) 4-40 x 3/8 FHMS, hyd end cap to inner plate | 55229 | 1 | 2.5 | 2.5 | 47.28 | 120.0912 | 300.228 | 0 | 0 | 47.28 | 120.0912 | 0 | 0 | 0 | 0 | 0
4 | Pulley, LeDuc pump | 49806 | 1 | 91.1 | 91.1 | 46.77 | 118.7958 | 10822.2974 | 0 | 0 | 46.77 | 118.7958 | 0 | 0 | 0 | 0 | 0
4 | Timing Belt | 52357 | 1 | 4 | 4 | 46.77 | 118.7958 | 475.1832 | 0 | 0 | 46.77 | 118.7958 | 0 | 0.5 | 1.27 | 5.08 | 0
4 | Pump Plate, Outer | 49805 | 1 | 96 | 96 | 46.5 | 118.11 | 11338.56 | 0 | 0 | 46.5 | 118.11 | 0 | 0.75 | 1.905 | 182.88 | 0
4 | (4x) shoulder screw 3/16 dia x 3/4 L, inner plate to outer plate |  | 1 | 17.3 | 17.3 | 46.98 | 119.3292 | 2064.3952 | 0 | 0 | 46.98 | 119.3292 | 0 | 0 | 0 | 0 | 0
4 | (4x) 8-32 nylok, inner plate to outer plate |  | 1 | 7.2 | 7.2 | 47.45 | 120.523 | 867.7656 | 0 | 0 | 47.45 | 120.523 | 0 | 0 | 0 | 0 | 0
3 | Assy., Boost Pump | 52378
| Subtotal Weight | 380.2 | gm | grams: diff
| Scale Weight | 380.6 | gm | grams: 0.4
4 | Pump Bracket | 49808 | 1 | 21.5 | 21.5 | 48.91 | 124.2314 | 2670.9751 | 0 | 0 | 48.91 | 124.2314 | 0 | 0 | 0 | 0 | 0
4 | Motor, Chrge Pump w/ W34 | 52288 | 1 | 60.6 | 60.6 | 48.9 | 124.206 | 7526.8836 | 0 | 0 | 48.9 | 124.206 | 0 | 1 | 2.54 | 153.924 | 0
4 | (2x) M2 x 6 FHMS, motor to bracket | 55203 | 1 | 0.2 | 0.2 | 48.9 | 124.206 | 24.8412 | 0 | 0 | 48.9 | 124.206 | 0 | 0 | 0 | 0 | 0
4 | Magnet Rotor Housing | 52457 | 1 | 10.6 | 10.6 | 48.9 | 124.206 | 1316.5836 | 0 | 0 | 48.9 | 124.206 | 0 | -0.4 | -1.016 | -10.7696 | 0
4 | (4x) 6-32 x 1/4 PHMS, bracket to housing | 55209 | 1 | 3.1 | 3.1 | 48.9 | 124.206 | 385.0386 | 0 | 0 | 48.9 | 124.206 | 0 | 0 | 0 | 0 | 0
4 | Rotary Magnet | 52455 | 1 | 18.1 | 18.1 | 48.9 | 124.206 | 2248.1286 | 0 | 0 | 48.9 | 124.206 | 0 | -0.4 | -1.016 | -18.3896 | 0
4 | Pumphead, Micropump + mounting plate + (3x) 4-40 x 3/8 FHMS | 52360 / 52456 / 55229 | 1 | 224.5 | 224.5 | 48.9 | 124.206 | 27884.247 | 0 | 0 | 48.9 | 124.206 | 0 | -0.8 | -2.032 | -456.184 | 0
4 | (4x) 4-40 x 3/16 SHCS + #4 LW, housing to flange | 55296 / 55233 | 1 | 1.9 | 1.9 | 48.9 | 124.206 | 235.9914 | 0 | 0 | 48.9 | 124.206 | 0 | -0.5 | -1.27 | -2.413 | 0
4 | (2x) Elbow, 1/8 npt x 1/4 tube, alum | 52431 | 1 | 25.2 | 25.2 | 48.9 | 124.206 | 3129.9912 | 0 | 0 | 48.9 | 124.206 | 0 | -0.9 | -2.286 | -57.6072 | 0
4 | Filter, Balston | 52425 | 1 | 14.5 | 14.5 | 47 | 119.38 | 1731.01 | 0 | 0 | 47 | 119.38 |  | 0 | 0 | 0 | 0
3 | Assy., Hydraulic Reservoir | 52346
| Subtotal Weight | 1112.3 | gm | grams: diff
| Scale Weight | 0 | gm | grams: -1112.3
4 | Piston, Diaphragm | 49811 | 1 | 172.7 | 172.7 | 45 | 114.3 | 19739.61 | 0 | 0 | 45 | 114.3 | 0 | 0 | 0 | 0 | 0
4 | Diaphragm, Internal (Bellofram) | 52359 | 1 | 104.2 | 596.9256 | 45.59 | 115.7985 | 69123.1182 | 0 | 0 | 45.59 | 115.7985 | 0 | 0 | 0 | 0 | 0
4 | Sikaflex Adhesive Compound | 55459 | 1 | 24.7 | 24.7 | 45 | 114.3 | 2823.21 | 0 | 0 | 45 | 114.3 | 0 | 0 | 0 | 0 | 0
4 | Cylinder Head, Diaphragm | 49813 | 1 | 326.5 | 326.5 | 43.9 | 111.506 | 36406.709 | 0 | 0 | 43.9 | 111.506 | 0 | 0 | 0 | 0 | 0
4 | Branch Tee, Male, 1/4 tube x 1/8 NPT x 1/4 tube, brass + Plastic plug (X59P4) | 52443 | 1 | 23.2 | 23.2 | 43.24 | 109.8296 | 2548.0467 | 0 | 0 | 43.24 | 109.8296 | 0 | -2.39 | -6.0706 | -140.8379 | 0
4 | Elbow, 1/8 NPT M x 1/8 NPT M, brass | 52426 | 1 | 20.5 | 20.5 | 43.24 | 109.8296 | 2251.5068 | 0 | 0 | 43.24 | 109.8296 | 0 | 2.39 | 6.0706 | 124.4473 | 0
4 | Kepner Chk Valve 404A-1-10 | 52447 | 1 | 25.5 | 25.5 | 43.24 | 109.8296 | 2800.6548 | 0 | 0 | 43.24 | 109.8296 | 0 | 0.8 | 2.032 | 51.816 | 0
4 | Elbow, 1/8 npt x 1/8 tube, alum | 52427 | 1 | 10.9 | 10.9 | 43.24 | 109.8296 | 1197.1426 | 0 | 0 | 43.24 | 109.8296 | 0 | -0.1 | -0.254 | -2.7686 | 0
4 | Cylinder, Diaphragm | 49812 | 1 | 326.5 | 326.5 | 45.3 | 115.062 | 37567.743 | 0 | 0 | 45.3 | 115.062 | 0 | 0 | 0 | 0 | 0
4 | (8x) 5-40 x 3/8 FHMS, head to cylinder | 55251 | 1 | 4.2 | 4.2 | 43.56 | 110.6424 | 464.6981 | 0 | 0 | 43.56 | 110.6424 | 0 | 0 | 0 | 0 | 0
4 | Standoff (starboard) | 55299 | 1 | 19.5 | 19.5 | 47.75 | 121.285 | 2365.0575 | 0 | 0 | 47.75 | 121.285 | 0 | 2.6 | 6.604 | 128.778 | 0
4 | Standoff (port) | 55299 | 1 | 19.5 | 19.5 | 47.75 | 121.285 | 2365.0575 | 0 | 0 | 47.75 | 121.285 | 0 | -2.6 | -6.604 | -128.778 | 0
4 | (4x) 6-32 x 1/2 SHCS, standoff to cylinder | 55183 | 1 | 4.9 | 4.9 | 47.5 | 120.65 | 591.185 | 0 | 0 | 47.5 | 120.65 | 0 | 0 | 0 | 0 | 0
4 | (2x) Saddles, mounting, Aft PCB + (4x) 4-40 x 1/4 PHMS | 52373 / 55257 | 1 | 6.4 | 6.4 | 45.5 | 115.57 | 739.648 | 0 | 0 | 45.5 | 115.57 | 0 | 3.2 | 8.128 | 52.0192 | 0
4 | (2x) Saddles, mounting, C PCB + (4x) 4-40 x 1/4 PHMS | 49888 / 55257 | 1 | 7.7 | 7.7 | 45.5 | 115.57 | 889.889 | 0 | 0 | 45.5 | 115.57 | 0 | 1.5 | 3.81 | 29.337 | 0
4 | (2x) Saddles, mounting, T PCB + (4x) 4-40 x 1/4 PHMS | 49888 / 55257 | 1 | 7.7 | 7.7 | 45.5 | 115.57 | 889.889 | 0 | 0 | 45.5 | 115.57 | 0 | -1.5 | -3.81 | -29.337 | 0
4 | (2x) Saddles, mounting, O2 PCB + (4x) 4-40 x 1/4 PHMS (only used with SBE 43f) | 49888 / 55257 | 1 | 7.7 | 7.7 | 45.5 | 115.57 | 889.889 | 0 | 0 | 45.5 | 115.57 | 0 | -3.2 | -8.128 | -62.5856 | 0
Weigh Hydraulic res. Assy, then Oil fill & add sub-assy fasteners
3 | Linear Poteniometer A w/ cable assy (W21), Linear pot clamp, 3-48 x 1/2 SHCS, Spring, Linear pot rod end | 52255 / 49814 / 55220 | 1 | 38.1 | 38.1 | 47.75 | 121.285 | 4620.9585 | 0 | 0 | 47.75 | 121.285 | 0 | 2.5 | 6.35 | 241.935 | 0
3 | Linear Poteniometer B w/ cable assy (W20), Linear pot clamp, 3-48 x 1/2 SHCS, Spring, Linear pot rod end | 52256 / 49814 / 55220 | 1 | 37.1 | 37.1 | 47.75 | 121.285 | 4499.6735 | 0 | 0 | 47.75 | 121.285 | 0 | 2.5 | 6.35 | 235.585 | 0
3 | (2x) 5-40 x 1/2 SHCS + #5 LW, Clamp to cylinder |  | 1 | 2.8 | 2.8 | 47.64 | 121.0056 | 338.8157 | 0 | 0 | 47.64 | 121.0056 | 0 | 0 | 0 | 0 | 0
3 | (2x) 6-32 x 1/2 SHCS + #6 LW, boost pump bracket to endcap | 55183 / 55217 | 1 | 2.7 | 2.7 | 48.28 | 122.6312 | 331.1042 | 0 | 0 | 48.28 | 122.6312 | 0 | 0 | 0 | 0 | 0
3 | (4x) 6-32 x 1/2 SHCS + #6 LW, cylinder standoff to endcap | 55183 / 55217 | 1 | 5.4 | 5.4 | 47.9 | 121.666 | 656.9964 | 0 | 0 | 47.9 | 121.666 | 0 | 0 | 0 | 0 | 0
aft endcap electronic parts - mount with CT sail
3 | Aft Terminal PCB | 55302 | 1 | 25.3 | 25.3 | 45.5 | 115.57 | 2923.921 | 0 | 0 | 45.5 | 115.57 | 0 | 3.4 | 8.636 | 218.4908 | 0
3 | SBE Conductivity PCB (included with sensors) | 55385 | 1 | 33.8 | 33.8 | 45.5 | 115.57 | 3906.266 | 0 | 0 | 45.5 | 115.57 | 0 | 1.7 | 4.318 | 145.9484 | 0
3 | Conductivity, PCB jumper, CA W26 | 52261 | 1 | 3.8 | 3.8 | 44.4 | 112.776 | 428.5488 | 0 | 0 | 44.4 | 112.776 | 0 | 1.5 | 3.81 | 14.478 | 0
3 | Shield, Radiation, C PCB + (4x) 2-56 x 1/4 BHCS | 49887 / 55221 | 1 | 20.9 | 20.9 | 45.5 | 115.57 | 2415.413 | 0 | 0 | 45.5 | 115.57 | 0 | 1.9 | 4.826 | 100.8634 | 0
3 | SBE Temperature PCB (included with sensors) | 55388 | 1 | 22.1 | 22.1 | 45.5 | 115.57 | 2554.097 | 0 | 0 | 45.5 | 115.57 | 0 | -1.7 | -4.318 | -95.4278 | 0
3 | Temperature, PCB jumper, CA W25 | 52260 | 1 | 4.6 | 4.6 | 44.4 | 112.776 | 518.7696 | 0 | 0 | 44.4 | 112.776 | 0 | 0.5 | 1.27 | 5.842 | 0
3 | Shield, Radiation, T PCB + (4x) 2-56 x 1/4 BHCS | 52371 / 55221 | 1 | 17.7 | 17.7 | 45.5 | 115.57 | 2045.589 | 0 | 0 | 45.5 | 115.57 | 0 | -1.9 | -4.826 | -85.4202 | 0
3 | SBE Oxygen PCB (only used with SBE 43f) ((included with sensors) | 55524 | 1 | 23.6 | 23.6 | 45.5 | 115.57 | 2727.452 | 0 | 0 | 45.5 | 115.57 | 0 | -3.4 | -8.636 | -203.8096 | 0
3 | Oxygen, PCB jumper, CA W33 (only used with SBE 43f) | 52265 | 1 | 5.5 | 5.5 | 44.4 | 112.776 | 620.268 | 0 | 0 | 44.4 | 112.776 | 0 | 0 | 0 | 0 | 0
3 | Ribbon Cable (W13) (gray) | 52248 | 1 | 30.5 | 30.5 | 38.625 | 98.1075 | 2992.2787 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.5 | 8.89 | 271.145 | 0
3 | Ribbon Cable (W14) (rainbow) | 52249 | 1 | 36.9 | 36.9 | 38.625 | 98.1075 | 3620.1667 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.5 | 8.89 | 328.041 | 0
3 | CT Sensor Sail Assembly (stanchion, guard, nut, nacelle, (6x) 6-32 x 5/16 FHMS, o-ring (E70-016), o-ring (E70-012), SBE components) | 52368 | 1 | 225.8 | 225.8 | 47.864 | 121.5746 | 27451.5356 | 184.027 | 184.027 | 47.864 | 121.5746 | 22373.0016 | 7.058 | 17.9273 | 4047.9889 | 3299.1109
2 | Assy., Mass shifter | 49889
| Subtotal Weight | 12494.2 | gm | grams: diff
| Scale Weight | 12485 | gm | grams: -9.2
3 | Assy., Bulkhead | 55843
4 | Hull Bulkhead | 49819 | 1 | 1307.3 | 1307.3 | 27.322 | 69.3979 | 90723.8485 | 2669.8 | 2669.8 | 27.506 | 69.8652 | 186526.2178 | 0 | 0 | 0 | 0
pressure test with partial fwd hull 49877 and complete 54844/5
3 | Centering Spring, Roll Bearing | 49845 | 1 | 8.5 | 8.5 | 27.58 | 70.0532 | 595.4522 | 0 | 0 | 27.58 | 70.0532 | 0 | 0 | 0 | 0 | 0
3 | Wire Bridge with vinyl tape wrap | 49894 | 1 | 9.8 | 9.8 | 26.8 | 68.072 | 667.1056 | 0 | 0 | 26.8 | 68.072 | 0 | 0 | 0 | 0 | 0
3 | (2x) 6-32 x 5/16 FHMS, wire bridge to bulkhead | 55208 | 1 | 3 | 3 | 26.8 | 68.072 | 204.216 | 0 | 0 | 26.8 | 68.072 | 0 | 0 | 0 | 0 | 0
3 | Snap Ring, Roll Bearing | 52402 | 1 | 53.8 | 53.8 | 27.8 | 70.612 | 3798.9256 | 0 | 0 | 27.8 | 70.612 | 0 | 0 | 0 | 0 | 0
3 | (4x) 6-32 x 1/2 FHMS, roll gear block to bulkhead | 55210 | 1 | 3.3 | 3.3 | 26.28 | 66.7512 | 220.279 | 0 | 0 | 26.28 | 66.7512 | 0 | -3.14 | -7.9756 | -26.3195 | 0
3 | Spur Gear, Delrin, Roll Pot. | 52468 | 1 | 6.8 | 6.8 | 27.25 | 69.215 | 470.662 | 0 | 0 | 27.25 | 69.215 | 0 | -2.43 | -6.1722 | -41.971 | 0
3 | Battery Pack, HV with CA W2 + shim | 52238 | 1 | 6516.2 | 6516.2 | 36.1166 | 91.7361 | 597770.7475 | 0 | 0 | 36.1166 | 91.7361 | 0 | -0.333 | -0.84582 | -5511.5323 | 0
3 | Bottom Ballast Weight | 52496 | 1 | 1149.5 | 1149.5 | 36.1166 | 91.7361 | 105450.6421 | 0 | 0 | 36.1166 | 91.7361 | 0 | -3.75 | -9.525 | -10948.9875 | 0
3 | (8x) 4-40 x 3/8 SHCS, ballast to batt encl endplates | 55228 | 1 | 4.8 | 4.8 | 36.1166 | 91.7361 | 440.3333 | 0 | 0 | 36.1166 | 91.7361 | 0 | -3.75 | -9.525 | -45.72 | 0
4 | Assy., Paine Pressure | 55484
| Subtotal Weight | 130.8 | gm | grams: diff
| Scale Weight | 131 | gm | grams: 0.2
5 | Paine Sensor w/ CA W12 | 52247 | 1 | 94.7 | 94.7 | 26.84 | 68.1736 | 6456.0399 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -120.269 | 0
5 | Male Connector, 1/16 NPT x 1/8 tube | 52438 | 1 | 15.2 | 15.2 | 26.84 | 68.1736 | 1036.2387 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -19.304 | 0
5 | Female Connector, 7/16-20 UNF x 1/8 tube + O-Ring (E70-012) | 49816 / 52409 | 1 | 18.7 | 18.7 | 26.84 | 68.1736 | 1274.8463 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -23.749 | 0
5 | Tubing, 0.125 D x 1.1 L SS | 52418 | 1 | 1.7 | 1.7 | 26.84 | 68.1736 | 115.8951 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -2.159 | 0
5 | FS-1265 Fluid in Pressure S. | 55472 | 1 | 0.5 | 0.5 | 26.84 | 68.1736 | 34.0868 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | -0.635 | 0
4 | Assy., Druck Pressure Sensor | 55485
| Subtotal Weight | 0 | gm | grams: diff
| Scale Weight | 0 | gm | grams: 0
5 | Druck Sensor w/ CA W12 | 55517 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Male Connector, 1/16 NPT x 1/8 tube | 52438 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Male connector, 1/4 NPT x 1/8 tube | 52347 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | Tubing, 0.125 D x 1.2 L SS | 52418 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
5 | FS-1265 Fluid in Pressure S. | 55472 | 1 | 0 | 0 | 26.84 | 68.1736 | 0 | 0 | 0 | 26.84 | 68.1736 | 0 | -0.5 | -1.27 | 0 | 0
3 | Assy., Pitch Mechanism | 55531
| Subtotal Weight | 2950 | gm | grams: diff
| Scale Weight | 2950 | gm | grams: 0
4 | (2x) Bracket, Linear Brng Rail-fwd | 49822 | 1 | 30.6 | 30.6 | 28.58 | 72.5932 | 2221.3519 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -169.4383 | 0
4 | (2x) Bracket, Linear Brng Rail-aft | 49822 | 1 | 30.4 | 30.4 | 41.72 | 105.9688 | 3221.4515 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -168.3309 | 0
4 | (4x) Linear Bearing Car | 55532 | 1 | 123 | 123 | 36.1166 | 91.7361 | 11283.5398 | 0 | 0 | 36.1166 | 91.7361 | 0 | -1.75 | -4.445 | -546.735 | 0
4 | (2x) Linear Brng Rail; THK RSR12VM | 52440 | 1 | 430.5 | 430.5 | 35.2 | 89.408 | 38490.144 | 0 | 0 | 35.2 | 89.408 | 0 | -2.18 | -5.5372 | -2383.7646 | 0
4 | (4x) 6-32 x 5/8 FHMS + 6-32 nylok, linear rail to bracket - fwd | 55211 / 55216 | 1 | 8.5 | 8.5 | 28.58 | 72.5932 | 617.0422 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -47.0662 | 0
4 | (4x) 6-32 x 5/8 FHMS + 6-32 nylok, linear rail to bracket - aft | 55211 / 55216 | 1 | 8.5 | 8.5 | 41.72 | 105.9688 | 900.7348 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -47.0662 | 0
4 | (2x) Batt. encl side plate | 52492 | 1 | 166.2 | 166.2 | 36.1166 | 91.7361 | 15246.5391 | 0 | 0 | 36.1166 | 91.7361 | 0 | -2.14 | -5.4356 | -903.3967 | 0
4 | (16x) M3 x 6 FHCS, batt. encl side plate to bearing car | 55202 | 1 | 6 | 6 | 36.1166 | 91.7361 | 550.4166 | 0 | 0 | 36.1166 | 91.7361 | 0 | -1.75 | -4.445 | -26.67 | 0
4 | Batt. encl endplate, fwd | 49869 | 1 | 112 | 112 | 32.1786 | 81.7336 | 9154.1605 | 0 | 0 | 32.1786 | 81.7336 | 0 | -0.33 | -0.8382 | -93.8784 | 0
4 | Batt. encl endplate, aft | 49869 | 1 | 111.9 | 111.9 | 40.0546 | 101.7386 | 11384.5511 | 0 | 0 | 40.0546 | 101.7386 | 0 | -0.33 | -0.8382 | -93.7946 | 0
4 | (8x) 4-40 x 3/8 FHCS, side plate to batt encl ends | 55229 / 55280 | 1 | 4 | 4 | 36.1166 | 91.7361 | 366.9444 | 0 | 0 | 36.1166 | 91.7361 | 0 | -2.14 | -5.4356 | -21.7424 | 0
4 | Fwd Endplate, MassShifter | 49824 | 1 | 278.7 | 278.7 | 27.63 | 70.1802 | 19559.2217 | 0 | 0 | 27.63 | 70.1802 | 0 | 0 | 0 | 0 | 0
4 | Bearing, Ball Screw Shaft-fwd, .375 ID x .875 OD | 52458 | 1 | 11.6 | 11.6 | 27.75 | 70.485 | 817.626 | 0 | 0 | 27.75 | 70.485 | 0 | 3.25 | 8.255 | 95.758 | 0
4 | (2x) Retaining Ring, internal spiral, Bearing-fwd | 55271 | 1 | 1.2 | 1.2 | 27.75 | 70.485 | 84.582 | 0 | 0 | 27.75 | 70.485 | 0 | 3.25 | 8.255 | 9.906 | 0
4 | Aft End Plate, MassShifter | 49821 | 1 | 93.5 | 93.5 | 42.9 | 108.966 | 10188.321 | 0 | 0 | 42.9 | 108.966 | 0 | 0 | 0 | 0 | 0
4 | Bearing, Ball Screw Shaft-aft, .375 ID x .875 OD | 52458 | 1 | 11.6 | 11.6 | 42.72 | 108.5088 | 1258.7021 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 95.758 | 0
4 | (2x) Retaining Ring, internal spiral, Bearing-aft | 55271 | 1 | 1.2 | 1.2 | 42.72 | 108.5088 | 130.2106 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 9.906 | 0
4 | Internal Ring Gear, Roll | 49854 / 52488 | 1 | 87.2 | 87.2 | 27.72 | 70.4088 | 6139.6474 | 0 | 0 | 27.72 | 70.4088 | 0 | -2.6 | -6.604 | -575.8688 | 0
4 | (6x) 4-40 x 3/8 FHMS, internal gear to fwd endplate | 55229 | 1 | 2.5 | 2.5 | 27.7 | 70.358 | 175.895 | 0 | 0 | 27.7 | 70.358 | 0 | -2.6 | -6.604 | -16.51 | 0
4 | Roll Bearing; KU075 XPO | 52487 | 1 | 411 | 411 | 27.57 | 70.0278 | 28781.4258 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | Ball Screw Shaft + spacers | 49820 / 55188 | 1 | 330.3 | 330.3 | 35.2 | 89.408 | 29531.4624 | 0 | 0 | 35.2 | 89.408 | 0 | 3.2 | 8.128 | 2684.6784 | 0
4 | Worm Gear, Pitch, Brass G1035 | 49853 | 1 | 48.6 | 48.6 | 27.15 | 68.961 | 3351.5046 | 0 | 0 | 27.15 | 68.961 | 0 | 3.23 | 8.2042 | 398.7241 | 0
4 | Roll Pin, pitch worm gear to shaft, 3/32 OD x 5/8 L | 55272 | 1 | 0.4 | 0.4 | 27.34 | 69.4436 | 27.7774 | 0 | 0 | 27.34 | 69.4436 | 0 | 3.225 | 8.1915 | 3.2766 | 0
4 | (3x) Retaining Plate; Roll Bearing | 49823 | 1 | 18.3 | 18.3 | 27.57 | 70.0278 | 1281.5087 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | (6x) 6-32 x 1/2 FHMS + (1x) 6-32 x 1/2 SHCS + #6 fender washer + #6 lock washer, ring plate to fwd plate |  | 1 | 8 | 8 | 27.57 | 70.0278 | 560.2224 | 0 | 0 | 27.57 | 70.0278 | 0 | 0 | 0 | 0 | 0
4 | Ball Screw Nut THK BTP0502 | 52403 | 1 | 215.7 | 215.7 | 36.1166 | 91.7361 | 19787.4759 | 0 | 0 | 36.1166 | 91.7361 | 0 | 3.2 | 8.128 | 1753.2096 | 0
4 | Mount, Ball Nut Screw | 52375 | 1 | 14.6 | 14.6 | 36.1166 | 91.7361 | 1339.347 | 0 | 0 | 36.1166 | 91.7361 | 0 | 2.9 | 7.366 | 107.5436 | 0
4 | (4x) 8-32 x 5/8 SHCS + #8 LW, bracket to batt end plate, fwd | 55244 / 55247 | 1 | 9.4 | 9.4 | 28.58 | 72.5932 | 682.3761 | 0 | 0 | 28.58 | 72.5932 | 0 | -2.18 | -5.5372 | -52.0497 | 0
4 | (4x) 8-32 x 5/8 SHCS + #8 LW, bracket to batt end plate, aft | 55244 / 55247 | 1 | 9.4 | 9.4 | 41.72 | 105.9688 | 996.1067 | 0 | 0 | 41.72 | 105.9688 | 0 | -2.18 | -5.5372 | -52.0497 | 0
4 | Retaining Ring, external .338 ID, Shaft-aft | 55263 | 1 | 0.1 | 0.1 | 42.72 | 108.5088 | 10.8509 | 0 | 0 | 42.72 | 108.5088 | 0 | 3.25 | 8.255 | 0.8255 | 0
4 | Top Plate | 52493 | 1 | 167.1 | 167.1 | 36.1166 | 91.7361 | 15329.1016 | 0 | 0 | 36.1166 | 91.7361 | 0 | 2.78 | 7.0612 | 1179.9265 | 0
4 | (4x) 4-40 x 3/8 SHCS + #4 FW, top plate assy to batt encl end | 55228 / 55280 | 1 | 2.7 | 2.7 | 36.1166 | 91.7361 | 247.6875 | 0 | 0 | 36.1166 | 91.7361 | 0 | -2.14 | -5.4356 | -14.6761 | 0
4 | (2x) 5/32 dia x 3/16 L SS SH Shoulder screw | 55213 / 55217 | 1 | 3.1 | 3.1 | 36.1166 | 91.7361 | 284.3819 | 0 | 0 | 36.1166 | 91.7361 | 0 | 2.9 | 7.366 | 22.8346 | 0
4 | Motor, Pitch with C.A. W41 | 52286 | 1 | 71.5 | 71.5 | 27.06 | 68.7324 | 4914.3666 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 426.7835 | 0
4 | Interface Plate, Pitch Motor | 49850 | 1 | 1.4 | 1.4 | 27.06 | 68.7324 | 96.2254 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 8.3566 | 0
4 | (2x) M2 x 6 CHMS + M2 LW, motor to interface plate | 55201 / 55200 | 1 | 0.4 | 0.4 | 27.06 | 68.7324 | 27.493 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 2.3876 | 0
4 | Gear Block, Pitch Motor | 49825 | 1 | 24.3 | 24.3 | 27.54 | 69.9516 | 1699.8239 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 147.5156 | 0
4 | (2x) Bearings, Pitch Motor, .25 ID x .625 OD flanged | 52452 | 1 | 9.2 | 9.2 | 27.06 | 68.7324 | 632.3381 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 54.9148 | 0
4 | Spacers, Pitch motor shaft | 55187 / 55267 / 55268 / 55260 | 1 | 2.7 | 2.7 | 27.06 | 68.7324 | 185.5775 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 16.1163 | 0
4 | Worm, Pitch | 52467 | 1 | 16 | 16 | 27.06 | 68.7324 | 1099.7184 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 95.504 | 0
4 | Roll pin, pitch worm to motor shaft, .094 OD x .375 L | 55276 | 1 | 0.2 | 0.2 | 27.06 | 68.7324 | 13.7465 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 1.1938 | 0
4 | (4x) 2-56 x 5/16 PHMS, interface plate to gear block | 55219 | 1 | 1 | 1 | 27.06 | 68.7324 | 68.7324 | 0 | 0 | 27.06 | 68.7324 | 0 | 2.35 | 5.969 | 5.969 | 0
4 | Gear Block shim plates, pitch motor | 55533 | 1 | 1.3 | 1.3 | 27.54 | 69.9516 | 90.9371 | 0 | 0 | 27.54 | 69.9516 |  | 2.39 | 6.0706 | 7.8918 | 0
4 | (4x) Shoulder screw, .125 dia x 7/16 L | 55222 | 1 | 6 | 6 | 27.54 | 69.9516 | 419.7096 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 36.4236 | 0
4 | (4x) # 4 nylok | 55232 | 1 | 2.1 | 2.1 | 27.54 | 69.9516 | 146.8984 | 0 | 0 | 27.54 | 69.9516 | 0 | 2.39 | 6.0706 | 12.7483 | 0
4 | 10 turn Potentiometer w/ nut and cable W41, pitch | 52287 | 1 | 33.6 | 33.6 | 27.72 | 70.4088 | 2365.7357 | 0 | 0 | 27.72 | 70.4088 | 0 | 1.99 | 5.0546 | 169.8346 | 0
4 | Spur gear, Pitch shaft, 48 tooth, delrin | 52460 | 1 | 9.1 | 9.1 | 26.8 | 68.072 | 619.4552 | 0 | 0 | 26.8 | 68.072 | 0 | 3.24 | 8.2296 | 74.8894 | 0
4 | Pitch pot gear, 72 tooth, delrin | 52461 | 1 | 13.4 | 13.4 | 26.78 | 68.0212 | 911.4841 | 0 | 0 | 26.78 | 68.0212 | 0 | 2.01 | 5.1054 | 68.4124 | 0
*
3 | Assy., Roll Motor | 55530
| Subtotal Weight | 350.4 | gm | grams: diff
| Scale Weight | 350.6 | gm | grams: 0.2
4 | Gear Block, Roll Motor | 49827 | 1 | 105.7 | 105.7 | 26.9 | 68.326 | 7222.0582 | 0 | 0 | 26.9 | 68.326 | 0 | -2.43 | -6.1722 | -652.4015 | 0
4 | (4x) Bearing, Roll mechanism, .25 ID x .625 OD flanged | 52452 | 1 | 18.2 | 18.2 | 26.8 | 68.072 | 1238.9104 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -69.8043 | 0
4 | Motor, Roll w/ shaft and C.A. W15 | 52250 | 1 | 71.5 | 71.5 | 26.8 | 68.072 | 4867.148 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -274.2311 | 0
4 | Interface Plate, Roll Motor | 49850 | 1 | 1.7 | 1.7 | 26.8 | 68.072 | 115.7224 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -6.5202 | 0
4 | (2x) M2 x 6 CHMS + M2 LW, motor to interface plate | 55201 / 55200 | 1 | 0.4 | 0.4 | 26.8 | 68.072 | 27.2288 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -1.5342 | 0
4 | Spacers, Roll Motor Shaft | 55266 / 55267 / 55260 | 1 | 2.1 | 2.1 | 26.8 | 68.072 | 142.9512 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -8.0543 | 0
4 | Worm, Roll | 52489 | 1 | 10 | 10 | 26.8 | 68.072 | 680.72 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -38.354 | 0
4 | (4x) 2-56 x 5/16 PHMS, interface plate to gear block | 55219 | 1 | 1 | 1 | 26.8 | 68.072 | 68.072 | 0 | 0 | 26.8 | 68.072 | 0 | -1.51 | -3.8354 | -3.8354 | 0
4 | Roll gear shaft | 49828 | 1 | 9.5 | 9.5 | 26.92 | 68.3768 | 649.5796 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -63.2206 | 0
4 | Worm Gear, Roll | 49851 | 1 | 87.5 | 87.5 | 26.86 | 68.2244 | 5969.635 | 0 | 0 | 26.86 | 68.2244 | 0 | -2.62 | -6.6548 | -582.295 | 0
4 | Roll Pin, Roll Gear Shaft, .063 OD x .75 L | 55275 | 1 | 0.1 | 0.1 | 26.92 | 68.3768 | 6.8377 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -0.66548 | 0
4 | Spacers, Ret. Rings, roll gear shaft | 55266 / 55267 / 55268 / 55266 | 1 | 1 | 1 | 26.92 | 68.3768 | 68.3768 | 0 | 0 | 26.92 | 68.3768 | 0 | -2.62 | -6.6548 | -6.6548 | 0
4 | Spur Gear, Roll | 49872 | 1 | 10.5 | 10.5 | 27.64 | 70.2056 | 737.1588 | 0 | 0 | 27.64 | 70.2056 | 0 | -2.66 | -6.7564 | -70.9422 | 0
4 | Roll Pin, Roll Spur Gear, .063 dia x .500 | 55274 | 1 | 0.1 | 0.1 | 27.64 | 70.2056 | 7.0206 | 0 | 0 | 27.64 | 70.2056 | 0 | -2.66 | -6.7564 | -0.67564 | 0
4 | Potentiometer, Roll w/ nut and CA W16 | 52251 | 1 | 31.1 | 31.1 | 27.25 | 69.215 | 2152.5865 | 0 | 0 | 27.25 | 69.215 | 0 | -2.43 | -6.1722 | -191.9554 | 0
2 | Assy., Electronics | 52347
| Subtotal Weight | 3266.6 | gm | grams: diff
| Scale Weight | 3263.3 | gm | grams: -3.3
3 | Electronics rail, port | 49846 | 1 | 75.1 | 75.1 | 23.507 | 59.7078 | 4484.0543 | 0 | 0 | 23.507 | 59.7078 | 0 | 0 | 0 | 0 | 0
3 | Electronics rail, starboard | 52379 | 1 | 76.8 | 76.8 | 23.507 | 59.7078 | 4585.5575 | 0 | 0 | 23.507 | 59.7078 | 0 | 0 | 0 | 0 | 0
3 | Seaglider Mainboard (Mother board, TT8, CF8, flash card) | 52290-B.4 | 1 | 269.8 | 269.8 | 21.5 | 54.61 | 14733.778 | 0 | 0 | 21.5 | 54.61 | 0 | 0.7 | 1.778 | 479.7044 | 0
3 | Insulation Sheet-OEM/Main | 52370 | 1 | 4 | 4 | 21.73 | 55.1942 | 220.7768 | 0 | 0 | 21.73 | 55.1942 | 0 | 0.44 | 1.1176 | 4.4704 | 0
3 | (6x) 6-32 x 1/4 PHMS, main board to rails | 55237 | 1 | 4.4 | 4.4 | 21.5 | 54.61 | 240.284 | 0 | 0 | 21.5 | 54.61 | 0 | 0.5 | 1.27 | 5.588 | 0
3 | Assy., OEM Navigation Board (transponder PCB, transformer/inductor, capacitors, GPS reciever, W6 GPS serial interface, GPS antenna cable,DC-DC converter, Cable Assembly W7, all soldered on wiring ) | 52257 | 1 | 588.7 | 588.7 | 21.5 | 54.61 | 32148.907 | 0 | 0 | 21.5 | 54.61 | 0 | -0.43 | -1.0922 | -642.9781 | 0
3 | (6x) 6-32 x 1/4 PHMS, OEM board to rails | 55237 | 1 | 4.4 | 4.4 | 21.48 | 54.5592 | 240.0605 | 0 | 0 | 21.48 | 54.5592 | 0 | -0.43 | -1.0922 | -4.8057 | 0
3 | Assy., Iridium Modem and bracket (includes satellite reciever - iridium LBT, cable assy W3, all clamps for wires, RF switch, DC block, and mounting hardware) | 52324 | 1 | 279.1 | 279.1 | 21.6 | 54.864 | 15312.5424 | 0 | 0 | 21.6 | 54.864 | 0 | 1.9 | 4.826 | 1346.9366 | 0
3 | (4x) 6-32 x 1/4 PHMS + #6 LW, modem cage to rails | 55209 / 55217 | 1 | 3.6 | 3.6 | 20.77 | 52.7558 | 189.9209 | 0 | 0 | 20.77 | 52.7558 | 0 | 0.29 | 0.7366 | 2.6518 | 0
3 | Battery Pack, LV w/ CA W1 | 52237 | 1 | 1881.5 | 1881.5 | 21.6 | 54.864 | 103226.616 | 0 | 0 | 21.6 | 54.864 | 0 | -1.81 | -4.5974 | -8650.0081 | 0
3 | Battery Pack Cage, LV | 49847 | 1 | 73.8 | 73.8 | 21.6 | 54.864 | 4048.9632 | 0 | 0 | 21.6 | 54.864 | 0 | -1.64 | -4.1656 | -307.4213 | 0
3 | (6x) 6-32 x 1/4 PHMS + #6 LW, batt. cage to rails | 55209 / 55217 | 1 | 5.4 | 5.4 | 21.78 | 55.3212 | 298.7345 | 0 | 0 | 21.78 | 55.3212 | 0 | -0.25 | -0.635 | -3.429 | 0
2 | Assy., Battery Hull Sections | 49878
| Subtotal Weight | 7983.6 | gm | grams: diff
| Scale Weight | 7976 | gm | grams: -7.6
3 | Hull, Fwd Battery | 49817 | 1 | 3614.7 | 3614.7 | 43.539 | 110.5891 | 399746.2752 | 12257.8 | 12257.8 | 43.539 | 110.5891 | 1355578.5797 | 0 | 0 | 0 | 0
3 | Hull, Aft Battery | 49817 | 1 | 3605 | 3605 | 33.671 | 85.5243 | 308315.2457 | 12257.8 | 12257.8 | 33.671 | 85.5243 | 1048340.2549 | 0 | 0 | 0 | 0
3 | Joint Ring , battery hulls | 49818 | 1 | 678.5 | 678.5 | 38.935 | 98.8949 | 67100.1897 | 436.8 | 436.8 | 38.935 | 98.8949 | 43197.2923 | 0 | 0 | 0 | 0
3 | (16x) 6-32 x 5/8 SHCS + #6 LW, joint ring to fwd/aft batt. Hull | 55214 / 55217 | 1 | 24 | 24 | 38.625 | 98.1075 | 2354.58 | 0 | 0 | 38.625 | 98.1075 | 0 | 0 | 0 | 0 | 0
3 | (2x) O-Ring, E70-271 | 52412 | 1 | 16.4 | 16.4 | 38.625 | 98.1075 | 1608.963 | 7.32 | 7.32 | 38.625 | 98.1075 | 718.1469 | 0 | 0 | 0 | 0
3 | (12x) Clips, Rib | 49883 | 1 | 20.9 | 20.9 | 38.625 | 98.1075 | 2050.4468 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 209.6897 | 0
3 | (12x) Clips, Wiring | 52405 | 1 | 18.6 | 18.6 | 38.625 | 98.1075 | 1824.7995 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 186.6138 | 0
3 | Oil Absorption pad, 1.9" x 10.125" + (2x) rib clips | 55527 / 49883 | 1 | 5.5 | 5.5 | 38.625 | 98.1075 | 539.5913 | 0 | 0 | 38.625 | 98.1075 | 0 | 3.95 | 10.033 | 55.1815 | 0
2 | Assy., Forward Hull Sections | 49877
| Subtotal Weight | 5150.9 | gm | grams: diff
| Scale Weight | 5152 | gm | grams: 1.1
3 | Endcap, Forward | 49830 | 1 | 1895.4 | 1895.4 | 15.566 | 39.5376 | 74939.6429 | 5178.3 | 5178.3 | 15.781 | 40.0837 | 207565.6308 | 0 | 0 | 0 | 0
3 | Acoustic transducer (ITC-3013), w/ cable W36 | 52267 / 52223 | 1 | 1439.9 | 1439.9 | 9.015 | 22.8981 | 32970.9742 | 656.2 | 656.2 | 9.015 | 22.8981 | 15025.7332 | 0 | 0 | 0 | 0
3 | O-Ring, E70-238, transducer to Endcap | 52413 | 1 | 3.2 | 3.2 | 10.5 | 26.67 | 85.344 | 0 | 0 | 10.5 | 26.67 | 0 | 0 | 0 | 0 | 0
3 | (6x) Alum 1/4-20 x 3/4 SHCS + 1/4 LW, xducer to end cap | 55236 / 55252 | 1 | 15.2 | 15.2 | 10.7 | 27.178 | 413.1056 | 0 | 0 | 10.7 | 27.178 | 0 | 0 | 0 | 0 | 0
3 | Hull, Electronics | 49829 | 1 | 1679.4 | 1679.4 | 22.843 | 58.0212 | 97440.8369 | 5240.9 | 5240.9 | 22.919 | 58.2143 | 305095.1152 | 0 | 0 | 0 | 0
3 | O-Ring, E70-263, elec hull to fwd endcap | 52410 | 1 | 6.6 | 6.6 | 19.77 | 50.2158 | 331.4243 | 5.76 | 5.76 | 19.77 | 50.2158 | 289.243 | 0 | 0 | 0 | 0
3 | (8x) 6-32 x 1/2 SHCS + #6 LW, elec hull to fwd endcap | 55183 / 55217 | 1 | 10.8 | 10.8 | 19.87 | 50.4698 | 545.0738 | 0 | 0 | 19.87 | 50.4698 | 0 | 0 | 0 | 0 | 0
(pressure test above components w/ bulkhead and pressure sensor assy)
3 | Compass bracket | 49848 | 1 | 24.2 | 24.2 | 12.01 | 30.5054 | 738.2307 | 0 | 0 | 12.01 | 30.5054 | 0 | -0.35 | -0.889 | -21.5138 | 0
3 | Compass w/compass carrier PCB | 55305 / 52466 | 1 | 45.8 | 45.8 | 12.01 | 30.5054 | 1397.1473 | 0 | 0 | 12.01 | 30.5054 | 0 | 0 | 0 | 0 | 0
3 | (4x) Nylon 4-40 x 3/8 PHMS + #4 x 1/8L stand off | 55256 / 55282 | 1 | 0.5 | 0.5 | 12.01 | 30.5054 | 15.2527 | 0 | 0 | 12.01 | 30.5054 | 0 | -0.31 | -0.7874 | -0.3937 | 0
3 | (2x) 4-40 X 3/8 SHCS + #4 LW, comp mt to fwd encap | 55228 / 55233 | 1 | 1.3 | 1.3 | 10.8 | 27.432 | 35.6616 | 0 | 0 | 10.8 | 27.432 | 0 | -0.81 | -2.0574 | -2.6746 | 0
3 | Cable Assembly W9 (compass cable) | 52244 | 1 | 28.6 | 28.6 | 15.33 | 38.9382 | 1113.6325 | 0 | 0 | 15.33 | 38.9382 | 0 | 0.38 | 0.9652 | 27.6047 | 0
Note on M2: Fritz Stahr:07 may 2006 Made changes to make easier and more clear to use. Added Ballast sheet with connections to both Trim sheet and Cal sheet. Added Pressure Sensor cal sheet directly here with connections to Cal sheet. Added Maintenance sheet to keep track of during build changes. Modified the Trim page to accomodate noted changes in weight of sub-assemblies. Changed top of page to interact with Ballast sheet better and to make explicit note of changes in trim weight location to get good pitch and roll gain. Russ' notes below are a good basic guide to the original intent of the sheet, but should not be taken too literally. Russ Light: (~2001?) There are two primary goals in achieve with the spreadsheet: 1) Trimming which involves getting neutral bouyancy 2) Balancing which involves achieving zero LCB-LCG separation Trimming the glider amounts to changing the weight value for the forward fairing trim weight until the Net Buoyancy value goes to zero. The Internal Oil Stroke must also be set for the desired percentage to achieve mission conditions. Thre stroke value determines the nuetral point of the VBD system and thus determines how much displacement is available for thrust and surface manuever. Of these two values the bouyancy required for the surface maneuver is the most important. The density of the surface water layer determines how much displacement in the external bladder will be required to achieve the proper antenna mast position for GPS and data telemetry. In cases of extream fresh water conditions as much as 650cc were required in the external bladder to achieve the proper surface position. Select the stroke percentage to give sufficient displacement for the surface maneuver. Adjust the forward fairing trim lead to zero the Net Buoyancy. To achieve the proper balance the LCB-LCG difference value is made to be zero. After achieving neutral buoyancy in the above procedure the center of gravity can be adjusted by changing the Pitch Mass Stroke percentage. The ideal value for this percentage is probably between 50-75%. The total throw of the pitch mass shifter is approx 15cm. Typical glider operations only require 1-3cm of movement to acheive normal glide slopes. By biasing the neutral point towards the aft more shift is provided for the surface manuever. FROM THE VMG USER'S GUIDE 10.0 Trimming Procedure This section outlines how to manipulate, or more generally, in what order to manipulate the VMG trim sheet. There is a specific trim sheet for every glider, and its maintenance and upkeep are a necessary and time-consuming task, especially with changes to the VM hardware. The goal is to set the VM Variable Buoyancy Device (VBD) for a given range of expected seawater density so that the limited range of buoyancy control is sufficient to perform its mission. After the vehicle is trimmed for neutral buoyancy at the expected density, it is then necessary to set the vehicle longitudinal pitch trim so that it floats level, with zero pitch, with sufficient remaining control over its pitch to dive and climb. The range of seawater density determines how precisely the VM must be trimmed. Density ranges greater than 0.7 practical salinity units (psu) may reduce the amount of thrust available during a dive to less than 200 grams. Such a density range will require exact laboratory trimming to avoid using some of the VBD volume for achieving neutral trim. The volume at the surface defines the minimum positive buoyancy and provides a starting point for the calculations. An estimate of the surface density (at an average depth of 0.5 to 1.0 meter) must be available for this calculation. With all the trim sheet weights and positions entered with utmost accuracy, we’re ready to set the final trim for a given density range expected. Adjustments are generally left to modifying the amount and location of external hull trim lead, which is that lead carried, or taped, to the inside of the bottom of the fairing. First we’ll set the overall buoyancy trim, then deal with the longitudinal pitch trim. Cell names in Italics represent actual cells in the spreadsheet. Set Overall Buoyancy Trim 1) Set the VBD Internal Oil Stroke to 5%. This represents the near maximum displaced volume, leaving 5% margin for trim errors and/or lighter than expected surface density. 2) Set Density to the expected surface density (e.g. 1.018) 3) Adjust the external trim lead to yield about 275 grams (+/- 10) grams of Net Buoyancy. This is required to elevate the present cellular phone/GPS antenna design to the maximum degree above the water. Future designs may take more or less: it is a function solely of the displaced volume of the antenna assembly. 4) Now set the Density to the expected seawater density at the apogee depth (e.g. 1.024). This is the density where the VM trim neutral point is defined. 5) Using the goal seek function of Microsoft EXCEL, or the hunt and peck method, determine the VBD internal oil gage setting to achieve neutral buoyancy at apogee. Force the Net Buoyancy value to zero by adjusting the value of the Internal Oil Stroke. Record this value of neutral trim in the calibration sheet (specifically the A/D value of VBD Position) 6) Observe the maximum obtainable negative buoyancy by setting the Internal Oil Stroke to 95%. This is the maximum thrust available during a dive. This should exceed that required for the mission as determined by expected currents or round trip dive times. If this maximum thrust is much more than is likely needed or desired, it might be wise to repeat the process, but leaving more positive buoyancy available for surfacing in lighter than expected surface water density (e.g. 10% or 15% versus 5%). Set Vehicle Longitudinal Pitch Trim 7) With the VBD set at the neutral trim point as determined in step #5, and the Pitch Mass Stroke set to 50%, play with the XCG position of the Trim Lead, Main Hull. Verify that there is position in the hull to accommodate the lead; this can be done in the big AutoCAD assembly drawing, #49800. It will probably be needed to place the trim lead in the most advantageous position in the fairing, and achieve the final pitch trim using the Pitch Mass Stroke, to achieve a zero Vehicle Pitch Angle. As done in step 5, record the Pitch Mass Position A/D value on the calibration sheet. Note that these two numbers, which define the neutral buoyancy and neutral pitch trim, are initial estimates only. Their values will surely be modified in the field based upon dive performance data.
Note on C18: This should be either measured or predicted denstiy the density of water the glider will encounter, e.g., OSB tank, PS surface, PS deep, WA coast deep.
Note on C19: Fritz (07May06): Minimum value for surface manuever is 150 cc so test by setting to density to lightest possible (e.g., 1.018 for PS) and then put oil reservoir to 5% to be sure it can do it.
Note on B20: Russ Light: The percentage of oil in the internal bladder. 100% stroke means the internal bladder is full. Since the spreadsheet is designed to be adjusted until neutral buoyancy is achieved this value is critical. The user must weigh the issues of how much positive buoyancy is required at the surface for surface maneuver based on surface water density. For fresh surface water conditions this can result in a lack of thrust range. Very fresh water conditions found during Puget Sound testing in Possession Sound during the spring/summer required surface buoyance values from 600-650 cc. In general the stoke is biased above 50% so that there is sufficient reserve for achieving enough positive buoyancy on the surface.
Note on E20: External Bladder Full = low AD = vehicle floats higher. External Bladder Empty = high AD = vehicle sinks lower.
Note on B27: Nominal is 70% so can pitch far forward during surface manuever, Max Stroke is 15.32 cm
Note on E27: Positive cm is aft = high AD. Negative cm is fwd = low AD.
Note on B35: Jim Osse: Itterates for roll by equating moment of battery against moment of hull VCG and hull VCB, must have itteration option ON
Note on J38: Fritz Stahr: Anything inside the pressure hull is considered to displace no volume.
Note on L38: Russ Light: In the case where the part is fabricated from a homogenous material the CG and CB are the same. In this case you will find that the CB is taken directly from the CG cell.
Note on F52: Fritz Stahr: determined by original all-up dry weight difference
Note on G52: Fritz Stahr: adjusted to make pitch stroke for tank or regression work
Note on J52: Fritz Stahr: set to make neutral observed in Port Susan show 0 net buoyancy
Note on L52: Fritz Stahr: adjusted to make pitch center found in PS be zero
Note on G62: Fritz Stahr: 16may06 In orig trim sheets this changed from glider to glider and appeared to be calculated. Further, this # was also pasted into the xcg for the aft fairing. Replaced with Solidworks locations for xcg/b.
Note on J62: Russ Light: Weighed fwd fairing (painted) with joint ring in air and in water: Air Wt = 6603g H20 Wt = 2595g Diff = Vol = 4008cc Using computed volumes of joint ring fwd fairing volume is = 3686cc
Note on G73: Mike Johnson: 42.077" represents center of mass for lead attached dirrectly to pupa. Lead is .377" further aft than lead taped to fairing
Note on G90: Tom: MLW solidworks centroid-no parts included 8/8/2001: 54.15" Jim: Orig ACAD estimate 35.20" total fairing; enter only for forward fairing Fritz: trim sheets ahd been taking this from fwd fairing, I assume to represent the whole fairing, but now individual LCGs and LCBs
Note on J91: mike johnson: 27 may 08 V= 695cc Volume taken from Solidworks model 49838-C (long wing)
Note on G102: Mike Johnson: 1/29/2008 LCG computed from solidworks model. Includes all components
Note on J102: mike johnson: 1/29/2008 Volume calculated with all components using solidworks models
Note on L102: mike johnson: 1/29/2008 Assume same as LCG
Note on G108: mike johnson: Distance for dual O2 aft fairing
Note on J108: mike johnson: 22 Jan 2008 Optode w/o cable = 81.1 cc Optode w/ potted cable = 105.22 cc Volume from solidworks part models
Note on J109: mike johnson: 22 Jan 2008 Cable = 29 cc Volume from solidworks part model
Note on G114: mike johnson: Distance for dual O2 aft fairing
Note on J114: mike johnson: 22 Jan 2008 SBE43f = 166.70 Volume from solidworks models of sensor, plenum, and bracket
Note on J115: mike johnson: 22 Jan 2008 24" IE55 cable = 29cc Volume from solidwork part model
Note on G121: mike johnson: Distance for dual WL aft fairing
Note on J121: Mike Johnson 22 Jan 2008 BB2F = 178.8cc Volume from solidworks part models of WL BB2F and (2x) sensor clamps
Note on G127: mike johnson: Distance for dual WL aft fairing
Note on J127: Mike Johnson 22 Jan 2008 WL VMG = 209.3 cc Volume from solidworks part models of WL vmg and (2x) sensor clamps
Note on J140: mike johnson: Calculated from OEM specs
Note on C159: mike johnson: Variable oil removed from weight roll up.
Note on J161: mike johnson: Volume taken from Solidworks model 49815-F. Internal features / thru holes removed to create a solid part.
Note on F187: Includes Oil
Note on J187: Jim Osse: incl vol of bladder at spec grav of 1.0
Note on B306: mike johnson: Higlighted are moving mass

## Lead Worksheet
SG lead worksheet | SG# | 114 |  | General notes re: usage
| date | 2008-06-23 |  | Lead, 1/8"x3/4"x5"≈ |  |  | 92g
| time | 7:30pm |  | Lead, 1/4"x3/4"x5"≈ |  |  | 170g
| who | FS
Add for ocean
| Wt. goal | Actual | [9 empty] | Actual
Forward Fairing | grams | grams/pcs | [9 empty] | sum grams
Nose plates (qty) | 1 | pc-1 | pc-2 | pc-3 | pc-4 | pc-5 | pc-6 | pc-7 | pc-8 | pc-9 | pc-10
weights | grams 1: 759; grams/pcs pc-1: 268.5; grams/pcs pc-2: 245.3; grams/pcs pc-3: 245.2 | 759
screws | grams 1: 23.7
Scale weight of fwd fairing with nose plates added
Scale weight of Pupa with lead attached:
Lead On Pupa At Fairing Joint Ring | grams/pcs pc-1: port/top; grams/pcs pc-5: middle; grams/pcs pc-9: stbrd/bottom
Bottom | 1452.8 |  | 177.6 | 170.5 | 170.6 | 182.5 | 182.4 | 175.6 | 170.4 | 168.2 | 55 | 1452.8
|  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.5 | 8.5
Port side | grams 1: 392.2; grams/pcs pc-3: 97.3; grams/pcs pc-4: 98.2; grams/pcs pc-5: 98.5; grams/pcs pc-6: 98.2 | 392.2
4 | bar: grams/pcs pc-3 to grams/pcs pc-6
Starboard side | grams 1: 487.1; grams/pcs pc-3: 95.5; grams/pcs pc-4: 96.7; grams/pcs pc-5: 98.4; grams/pcs pc-6: 98; grams/pcs pc-7: 98.5 | 487.1
5 | bar: grams/pcs pc-3 to grams/pcs pc-7
Top | grams 1: 454.9; grams/pcs pc-3: 97.5; grams/pcs pc-4: 98; grams/pcs pc-5: 98.7; grams/pcs pc-6: 96.5; grams/pcs pc-7: 64.2 | 454.9
| grams/pcs pc-3: 1; grams/pcs pc-4: 1; grams/pcs pc-5: 1; grams/pcs pc-6: 1; grams/pcs pc-7: 0.5 | 4.5
sum - rubber | grams 1: 218.2 | 218.2
Sum tape | [11 empty] | 22
Lead On Pupa At Electronics Hull | grams 1: 0 | 0
sum - rubber | grams 1: 0 | 0
Aft Fairing
At fwd edge stern tube | grams 1: 0 | 0
# of pcs tape | grams 1: 0
Scale weight of aft fairing after lead added
Whole glider weight - BEFORE PUTTING IN TANK - DRY!!!

## Ballast
Ballasting worksheet |  | SG | 114
| Date: | 2008-06-24 10:57 | Time: | 39623.4559
| Conversions |  | Units
VBD | -4.0767 | counts/cc | cc=cm3 | gm=grams
| -0.2453 | cc/counts | counts=AtoD counts
In tank to find original volume of SG
Tank Temp | 23.93 | deg C | measured using YSI
Tank Salinity (meas) | 35.04 | ppt | measured using YSI
Tank Salinity (corr) | 33.8682 | ppt | 0.96656 | ratio for tank salt
Tank Density | 1.0228 | gram/cc | calculated using web-calculator
SG Mass | 52408 | grams | measured with lead est. from orig. trim sheet
SG Volume | 51238.7321 | cc | at neutral, when achieved relative to C_VBD
Neutral by interation in tank |  | 4.4 | cc
Relative to set C_VBD |  | 1908 | counts
(Neil's table ) | counts* | VBD rel V0, cc | Displacement
Vol max (oil outside) | 420 | 360.6004 | 51599.3326 | cc (pos. floats)
Vol neutral | 1890.0625 | (obsv in tank) | 51238.7321 | cc
Vol min (oil inside) | 3949 | -505.0491 | 50733.683 | cc (neg. sinks)
Max movable vol. |  | 865.6496
New enviornment ballasting
SG Mass (observed) | 51974 | gm | at PS deployment 19 Jun 08
SG Vol max (obsv/est) | 51568 | cc | from PS deployment regression 23 Jun 08
SG Vol min | 50702.3504 | cc
New enivron density | 1.0275 | g/cc
Goal for thrust | 300 | cc
New Vol neutral | 51002.3504 | cc | lead density | 11.296 | gm/cc
Change mass by | 474.0338 | gm | vol lead change | 41.9648 | cc
Projected new mass | 52448.0338 | gm
Proj. Vol. max | 51609.9648 | cc
VBD points for new environ
(Neil's table ) | counts* | VBD rel V0, cc | Displacement
Vol max | 420 | 565.6496 | 51609.9648 | cc
Vol neutral (C_VBD) | 2725.9878 | 0 | 51002.3504 | cc
Vol min | 3949 | -300 | 50744.3152 | cc
Max movable vol. |  | 865.6496
Actual new mass | 52446.2 | gm
Est. Vol. neutral | 51042.5304 | cc
Est. C_VBD | 2562.1857 | counts

## Tank _ PS notes
Tank and Puget Sound Notes
Seaglider | 114
Date | Action
Weighed today > overnight drying - total 51794 g. Put in tank, found neutral at about +28cc with cable on relative to C_VBD of 3339. So why won't this thing sink in Port Susan?!? - check log and cap files carefullyto see what's up!
Discovered that the connector for the pressure transducer was off by one pin, so appeared to work at sealevel, but didn't work when in water. Also changed out transponder gboard
2006-06-16 | Flew well in PS yesterday and today. Neil did regressions over dives 2-19 and found that C_PITCH should have been 2732 to make flat and PITCH_GAIN should be 15. Use these to tune spreadsheet.
With mass = 51672g, Vol.max = 51233cc (starting point for regression). Dives 2-19 regressed to show true Vol.max=51233cc. So, if want 250 thrust and Vol.neutral in 27.5 sigma-t water, then add 451g lead. New Vol.max should be 51273, so set C_VBD to 2871. See Ballast worksheet for more #s.
2006-08-04 | Sent to PS without any tank work - ballasted close to previous PS runs - all up weight now 51737gm, set C_VBD to 3485 based on trim sheet
2006-08-08 | Neil did regressions over dives 3-23 in Port Susan and gets a vol-max of 51329
2007-04-09 | In tank with new weights prior to Port Susan testing to verify all good with new compass and CT calibration, and new springs on VBD. Removed ~550gm, then put in at 51625 gm, with density
took out 50gm, then re-checked in tank - found Cpitch better at 2784, CVBD for PS to be
4/11-12/2007 | In PS found better pitch at 2818 and better roll_dive at 2393 and roll_climb at 2362. These are rough estimates but put into spread sheet

## Maint
Maintenance sheet
SG | 114
Date | Action
2006-06-09 | changed CT sensor set due to noisy T. Now has SN 0038, while reassembling, broke transducer connector, repaired (but didn't check functionality!)
2006-06-13 | changed OEM board to SFC-OEM-022 with AAE SN 010-449, Garmin SN 27701784, DC-DC SN 01094A517A
2006-08-01 | Warranty work: replaced pitch pot, updated main.run, loaded 3.05 into AAE transponder board, reset pich/roll software limits,
added shielding to CT wires/PCBs (needed to unsolder/solder wires to implement changes), repaired short on main PCB (U67),
added epoxy paint over external scratches, removed material from aft endplate in mass shifter assy for clearance over reservoir check valve
2006-08-03 | Replace modem due to not getting good data uploads with any combination of antennas, SIMs, etc.
2007-03-15 | Glider returned to SFC. Weight placement in fwd fairing did not match trim sheet (wts in glider: top 612.2, starboard 338.9, bottom 1244.0, port 392.4, fwd bottom 532.7). Looked like some weight had come loose and had been retaped. All weight removed and retaped according to trim sheet. Trim sheet updated to reflect retaped weight (see SG-114 hawaii 8aug06 for previous trim sheet).
2007-03-19 | Oil in aft endcap. Looked like bellowfram leaking. Replaced bellowfram/piston (piston new short version). Added springs (model MC103-0905-S w/washers) to linear pots. Re-set VBD limits.
2007-03-20 | Checked number of root file entries on CF card, = 4096 OK
2007-03-23 | Updated code to 1036M
2008-06-23 | Refrub at SFC (dec07-june08):
Replaced fwd fairing, aft fairing, wings, rudder, and antenna due to damage sustained in the beach surf
CT sail recalibrated
SBE 43f Oxygen sensor recalibrated, replaced O2 cable
Original Wetlabs BB2F damaged, replaced with new BB2F
Updated VBD system with new linpot clamps, springs, and rod ends. Re-set VBD limits
Rebuilt mass shifter: new 24V batt, new top plate w/shoulder screw, replaced linear rails/bearing cars and lubed with high slip grease, lubed ball screw with high slip grease, lubed pitch/roll gears with Bostik Blue moly, re-set pitch/roll limits, ran/passed 100 cycle pitch current test
Replaced 10V batt
Repaired burnt traces on main board
Recalibrated compass
Replaced modem due to poor comms with the new antenna
Updated trim sheet with new part/batt weights
Updated cal sheet with new calibration values
Ballasted glider with new lead on pupa technique
Ran/passed self test/deck dive
Flew 24 hour Port Susan check out dive
Re-ballasted for opean ocean

## Pres cal
Pressure Sensor |  | last updated | 2008-03-11
SG | 114
Manufacturer | Paine | Druck |  |  | Paine calibration data from sheet delivered with sensor
Model # | 211-75-710-05 | PCDR 4020 | Conversion to SG counts |  | Excitation Voltage = 10.000V
Serial # | 208104 | xxxxxx | A/D Ref Voltage | 2.495 | Response voltages in mV
Calibration date | 2003-06-10 | xxxxxx | Volts/AD-count | 1.4871e-07 | Pressure | 0 deg C |  | 30 deg C
A/D Gain | 128 | 64 | A/D Counts/Volt | 6724335.0701 | (PSIA) | Increasing | Decreasing | Increasing | Decreasing
Atmos @ cal. | 14.7 | 14.x | Max counts (24 bit AD) | 16777216 | 0 | -0.145 | -0.133 | -0.144 | -0.129
Temp @ cal. | 0 | xx.x | Excitation Voltage | 5.000 | 300 | 5.861 | 5.895 | 5.866 | 5.905
[5 empty] | 600 | 11.871 | 11.903 | 11.883 | 11.913
[5 empty] | 900 | 17.869 | 17.895 | 17.882 | 17.912
Druck calibration on APL dead-weight tester |  |  |  |  | 1200 | 23.845 | 23.861 | 23.863 | 23.874
Pressure (psig) | AD counts | Variance | Notes |  | 1500 | 29.807 |  | 29.821
From dead-weight | Mean reading | (using 40 samples)
0 |  |  | After setting y-intercept with 40 samples |  | Averages of increasing and decreasing and adjusted for Seaglider 5.000V excitation voltage
100 |  |  |  |  | (PSIA) | Average |  | Average
200 |  |  |  |  | 0 | -0.0695 |  | -0.06825
400 |  |  |  |  | 300 | 2.939 |  | 2.9428
600 |  |  |  |  | 600 | 5.9435 |  | 5.949
800 |  |  |  |  | 900 | 8.941 |  | 8.9485
1000 |  |  |  |  | 1200 | 11.9265 |  | 11.9342
1200 |  |  |  |  | 1500 | 14.9035 |  | 14.9105
1400
1600
1500 |  |  |  |  | 0 Deg C |  | 30 Deg C
1400 |  |  |  |  | Pressure | Output | Pressure | Output
1200 |  |  |  |  | (PSIA) | (A/D counts) | (PSIA) | (A/D counts)
1000 |  |  |  |  | 0 | -59819.6848 | 0 | -58743.7912
800 |  |  |  |  | 300 | 2529641.0587 | 300 | 2532868.7395
600 |  |  |  |  | 600 | 5115658.9426 | 600 | 5120392.8745
400 |  |  |  |  | 900 | 7695651.8224 | 900 | 7702107.184
200 |  |  |  |  | 1200 | 10265316.1234 | 1200 | 10271986.6638
100 |  |  |  |  | 1500 | 12827664.3479 | 1500 | 12833689.3521
0
Slope Coefficient Units (PSIA/A-D )
Slope calculations |  |  |  |  | Intercept Coefficient (PSIA)
Averages | Pressure (psia) | AD counts | Std deviation |  | 0 Deg C |  | 30 Deg C
| 0 | #DIV/0! | #DIV/0! |  | Slope | 0.00011638 | Slope | 0.00011633
| 100 | #DIV/0! | #DIV/0! |  | Intercept | 5.6959 | Intercept | 5.4535
| 200 | #DIV/0! | #DIV/0!
| 400 | #DIV/0! | #DIV/0!
| 600 | #DIV/0! | #DIV/0! |  | Error Analysis 0 Deg |  |  | Error Analysis 30 Deg
| 800 | #DIV/0! | #DIV/0! |  | Actual | Calculated | Error | Actual | Calculated | Error
| 1000 | #DIV/0! | #DIV/0! |  | (PSIA) | (PSIA) |  | (PSIA) | (PSIA)
| 1200 | #DIV/0! | #DIV/0! |  | 0 | -1.2656 |  | 0 | -1.3801
| 1400 | #DIV/0! | #DIV/0! |  | 300 | 300.0854 | -0.02845% | 300 | 300.0981 | -0.0327%
| 1500 | #DIV/0! |  |  | 600 | 601.0357 | -0.1726% | 600 | 601.1007 | -0.1835%
[5 empty] | 900 | 901.2848 | -0.1428% | 900 | 901.4275 | -0.1586%
Slope | #DIV/0! |  |  |  | 1200 | 1200.332 | -0.02767% | 1200 | 1200.3775 | -0.03146%
Intercept – psia | #DIV/0! |  |  |  | 1500 | 1498.5277 | 0.09815% | 1500 | 1498.3763 | 0.1082%
Intercept – psig | #DIV/0!
Use these values for Seaglider Calibration Coefficients
Slope Coefficient Units (PSIG/A-D )
Intercept Coefficient (PSIG)
[5 empty] | 0 Deg C |  |  | 30 Deg C
[5 empty] | Slope | 0.00011638 |  | Slope | 0.00011633
[5 empty] | Intercept | -9.0041 |  | Intercept | -14.7
[5 empty] | Depth Offset | -6.1678 |  | Depth Offset | -10.0695

## Motor Data
Motor Data
| Main Pump Motor |  |  |  | Boost Pump motor |  |  |  | Roll Motor |  |  |  | Pitch Motor |  |  |  | Amp w/full mass shifter assy. orientation - nose down, full lift
SG SN | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/main pump and oil (no load on bench) (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/micro pump (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/gear block assy. (mA): | Burn in time (min): | Amp @ start of burn in (mA): | Amp @ end of burn in (mA): | Amp w/gear block assy. (mA): | 100-cycle average time (sec): | 100-cycle avg. current (mA) | 100-cycle avg. current peak (mA)
114 | 30 | na | 31 | 515 | 30 | na | 10 | 25 | 30 | na | 10 | 26 | 30 | na | 23 | 35 | 22 | 189 | 242

## Material
Material Density | SG | 114
Date: | 2008-06-24 10:57 | Time: | 39623.4559
6061 T6 Alum | 2.713 | gm/cc
Lead | 11.296 | gm/cc
Naval Brass | 8.415 | gm/cc
Stainless Steel | 8.027 | gm/cc
Polypropylene | 0.9 | gm/cc
Fairing fiberglass density** | 1.6331 | gm/cc | from Port Susan 19 Jun 08
Zinc | 7.1 | gm/cc
PVC Density | 1.41 | gm/cc
Neoprene Density | 1.23 | gm/cc
Hydraulic Oil* Density | 0.8683 | gm/cc
PRC Wing Density | 0.672 | gm/cc
EPDM sheet density | 1.3936 | gm/cc
Nylon density | 1.151 | gm/cc
Polycarbonate | 1.21 | gm/cc
** vary to balance overall density of glider
*Mobil DTE 11M specfic gravity is 0.8683gm/cc weighed by Lehman 6/26/2001
Note on B8: Russ Light (~2000?): Using an in air and in water weight for aft fairing calculated the volume. Then calculated the density of 1.5384 from this volume and in air weight. (2006) Jim O finds this varies between ~ 1.45 & ~1.65 from DG work. Osse note 1/15/07; suggest using 1.43 as startign density for fiberglass based on average of SG 101 to SG 110 post field deployment data
Note on B13: Jim Osse: 0.672 gm/cc per Lehman test on 5/28/02 on wing#23

## Matlab
% sg_calib_constants.m
% values from spreadsheet for easy inclusion of glider params into diveplot.m, etc.
% basic glider and mission params
id_str | = | 114 | ;
mass | = | 52448.2 | ; | % from sum on trim sheet
volmax | = | 51609.9648 | ; | % from ballast worksheet final table
mission_title | = | SFC_Port_Susan | ;
rho0 | = | 1027.5 | ; | % from ballast worksheet for "new" evironment
% software limits from cal sheet:
pitch_min_cnts | = | 410 | ;
pitch_max_cnts | = | 3699 | ;
roll_min_cnts | = | 160 | ;
roll_max_cnts | = | 3940 | ;
vbd_min_cnts | = | 420 | ;
vbd_max_cnts | = | 3949 | ;
vbd_cnts_per_cc | = | -4.0767 | ;
% CT sensors cal constants
calibcomm | = | SN_0041_cal_25Apr06 | ; | % SN and cal date
t_g | = | 0.0044131 | ;
t_h | = | 0.00065298 | ;
t_i | = | 2.7242e-05 | ;
t_j | = | 2.9467e-06 | ;
c_g | = | -10.0024 | ;
c_h | = | 1.1227 | ;
c_i | = | -0.001173 | ;
c_j | = | 0.00011679 | ;
cpcor | = | -9.57e-08 | ;
ctcor | = | 3.25e-06 | ;
sbe_cond_freq_min | = | 2.9832 | ; | % kHz, from cal for 0 salinity
sbe_cond_freq_max | = | 8 | ; | % kHz, est for greater than 34.9 sal max T
sbe_temp_freq_min | = | 2.9242 | ; | % kHz, from cal for 1 deg T
sbe_temp_freq_max | = | 5.5553 | ; | % kHz, from cal for 32.5 deg T
% oxygen cal constants
comm_oxy_type | = | SBE_43f | ; | % spec SBE_43f or Aa_optode
calibcomm_oxygen | = | SN_430102_cal_25Apr06 | ; | % SN and cal date
Soc | = | 0 | ;
Foffset | = | -0.0017638 | ;
A | = | 0.00017276 | ;
B | = | -3.4856e-06 | ;
C | = | 0.036 | ;
D | = | #REF! | ;
% initial hydrodynamic model params
hd_a | = | 0.003836 | ; | % #s from SG115 in HI, July 2006
hd_b | = | 0.010078 | ;
hd_c | = | 9.8541e-06 | ;